## Measurements
Measured distance (m) | Measured illuminance (lx) | Calculated lum. intensity (cd)
1 | 213 | 213
1.5 | 94.9 | 213.525
2 | 53.4 | 213.6
2.5 | 34.2 | 213.75
3 | 23.7 | 213.3
| Mean: | 213.435
| Std Dev: | 0.26153
| 2SD/Avg: | 0.2451%

## Uncertainty types
Uncertainty component | Value | Distribution
Detector calibration | 1 % (k = 2) | Normal
Detector resolution | Variable | Rectangular
Distance meter calibration | 2 mm (k = 2) | Normal
Distance meter resolution | 0.5 mm | Rectangular

## Sampling
Simple example with individual samples (not exhaustive) - press F9 to refresh
| Meas | Uncertainties |  | Samples |  | Sample | Meas | Uncertainties |  | Samples |  | Sample | Calc
Meas | Illum | Det Cal | Det Res | Det Cal | Det Res | Illum | Dist | Meter Cal | Meter Res | Meter Cal | Meter Res | Dist | Int
No. | (lx) | (%) | (lx) | (lx) | (lx) | (lx) | (m) | (m) | (m) | (m) | (m) | (m) | (cd)
1 | 213 | 1 | 0.5 | -0.92487 | 0.44231 | 212.5174 | 1 | 0.002 | 0.0005 | 0.00019254 | 0.00048545 | 1.0007 | 212.8057
2 | 94.9 | 1 | 0.05 | 0.065567 | 0.04961 | 95.0152 | 1.5 | 0.002 | 0.0005 | 0.0019065 | -0.00014139 | 1.5018 | 214.2876
3 | 53.4 | 1 | 0.05 | -0.12629 | -0.044338 | 53.2294 | 2 | 0.002 | 0.0005 | -0.0011373 | -0.00011697 | 1.9987 | 212.6505
4 | 34.2 | 1 | 0.05 | -0.061647 | 0.0099136 | 34.1483 | 2.5 | 0.002 | 0.0005 | -0.001333 | -0.00019511 | 2.4985 | 213.1658
5 | 23.7 | 1 | 0.05 | 0.091926 | -0.0066804 | 23.7852 | 3 | 0.002 | 0.0005 | 7.3125e-05 | 0.00035842 | 3.0004 | 214.1288
[12 empty] | Average: | 213.4077

## Monte Carlo
[11 empty] | Absolute Uncert. |  | Relative Uncert.
Input parameters |  |  | Press F9 to refresh
Measured illuminance: | 23.7 | lx | [5 empty] | Original measured value: | 213.3 | cd
Calib uncertainty: | 1% | [6 empty] | Lowest 2.5% point: | 211.2685 | cd | -2.0315 | cd | -0.9524%
Resolution: | 0.05 | lx | Lower | 204.8898 |  |  |  | Highest 2.5% point: | 215.4585 | cd | 2.1585 | cd | 1.012%
Distance measured: | 3 | m | Upper | 221.442 |  |  |  | 2 standard deviations: | 2.1858 | cd
Distance meter cal: | 0.002 | m | Bins | 30 |  |  |  | 2 standard deviations: | 1.025%
Distance meter resolution: | 0.0005 | m | Bin width | 0.55174
5000
Samples (cd) |  | Sorted (cd) | # | Lower | Upper | Range | Count
215.1877 |  | 209.0713 | 0 | 204.8898 | 205.4416 | 204.9-205.4 | 0
215.0254 |  | 209.4289 | 1 | 205.4416 | 205.9933 | 205.4-206.0 | 0
212.9306 |  | 209.5057 | 2 | 205.9933 | 206.5451 | 206.0-206.5 | 0
214.2233 |  | 209.7003 | 3 | 206.5451 | 207.0968 | 206.5-207.1 | 0
212.8286 |  | 209.7004 | 4 | 207.0968 | 207.6485 | 207.1-207.6 | 0
212.5218 |  | 209.8875 | 5 | 207.6485 | 208.2003 | 207.6-208.2 | 0
211.8469 |  | 209.995 | 6 | 208.2003 | 208.752 | 208.2-208.8 | 0
213.2667 |  | 210.0217 | 7 | 208.752 | 209.3037 | 208.8-209.3 | 1
214.2469 |  | 210.1315 | 8 | 209.3037 | 209.8555 | 209.3-209.9 | 4
213.9474 |  | 210.2318 | 9 | 209.8555 | 210.4072 | 209.9-210.4 | 8
213.0832 |  | 210.2999 | 10 | 210.4072 | 210.959 | 210.4-211.0 | 52
214.1988 |  | 210.3348 | 11 | 210.959 | 211.5107 | 211.0-211.5 | 171
211.9815 |  | 210.3757 | 12 | 211.5107 | 212.0624 | 211.5-212.1 | 410
214.5968 |  | 210.4072 | 13 | 212.0624 | 212.6142 | 212.1-212.6 | 699
211.5444 |  | 210.432 | 14 | 212.6142 | 213.1659 | 212.6-213.2 | 929
212.424 |  | 210.4392 | 15 | 213.1659 | 213.7176 | 213.2-213.7 | 948
212.9671 |  | 210.484 | 16 | 213.7176 | 214.2694 | 213.7-214.3 | 810
214.1271 |  | 210.4993 | 17 | 214.2694 | 214.8211 | 214.3-214.8 | 561
214.101 |  | 210.5025 | 18 | 214.8211 | 215.3729 | 214.8-215.4 | 255
213.8118 |  | 210.5244 | 19 | 215.3729 | 215.9246 | 215.4-215.9 | 112
212.0296 |  | 210.5495 | 20 | 215.9246 | 216.4763 | 215.9-216.5 | 27
211.4378 |  | 210.5647 | 21 | 216.4763 | 217.0281 | 216.5-217.0 | 12
214.2515 |  | 210.5683 | 22 | 217.0281 | 217.5798 | 217.0-217.6 | 1
214.8236 |  | 210.5734 | 23 | 217.5798 | 218.1315 | 217.6-218.1 | 0
214.4539 |  | 210.5978 | 24 | 218.1315 | 218.6833 | 218.1-218.7 | 0
212.5612 |  | 210.6057 | 25 | 218.6833 | 219.235 | 218.7-219.2 | 0
211.9638 |  | 210.6108 | 26 | 219.235 | 219.7868 | 219.2-219.8 | 0
214.5781 |  | 210.6214 | 27 | 219.7868 | 220.3385 | 219.8-220.3 | 0
214.0725 |  | 210.6229 | 28 | 220.3385 | 220.8902 | 220.3-220.9 | 0
213.033 |  | 210.6238 | 29 | 220.8902 | 221.442 | 220.9-221.4 | 0
212.5504 |  | 210.6288
215.3095 |  | 210.6363
214.0728 |  | 210.6366
214.7224 |  | 210.6423
214.4621 |  | 210.6478
214.1914 |  | 210.6731
212.6771 |  | 210.6857
213.2815 |  | 210.6874
212.5139 |  | 210.6897
212.3199 |  | 210.6997
211.4404 |  | 210.7055
212.8827 |  | 210.7163
214.4599 |  | 210.7592
213.3507 |  | 210.7916
212.9098 |  | 210.7991
213.9202 |  | 210.8017
213.5095 |  | 210.8082
213.7683 |  | 210.8197
212.1673 |  | 210.8205
214.1874 |  | 210.8235
214.2752 |  | 210.8258
213.0479 |  | 210.8363
213.4036 |  | 210.8393
213.5143 |  | 210.8522
213.9964 |  | 210.8673
212.9194 |  | 210.8689
211.9945 |  | 210.8799
211.3519 |  | 210.8992
213.0404 |  | 210.8998
215.6079 |  | 210.9027
213.3487 |  | 210.9089
213.7389 |  | 210.9256
213.8768 |  | 210.9352
213.7931 |  | 210.9366
212.3365 |  | 210.9489
213.2146 |  | 210.9782
213.0694 |  | 210.9825
212.2483 |  | 210.9885
214.2265 |  | 210.9924
212.7306 |  | 210.9977
213.9408 |  | 211.0018
215.2546 |  | 211.005
212.158 |  | 211.0083
214.3877 |  | 211.0242
215.5494 |  | 211.0322
214.9341 |  | 211.0384
212.9179 |  | 211.0402
212.9327 |  | 211.0471
213.7428 |  | 211.0559
211.436 |  | 211.064
212.9198 |  | 211.0786
212.4901 |  | 211.0795
211.3452 |  | 211.0812
212.3256 |  | 211.096
214.874 |  | 211.1092
212.218 |  | 211.1094
213.4718 |  | 211.1124
216.1345 |  | 211.1193
212.1628 |  | 211.1221
215.3311 |  | 211.1223
212.7919 |  | 211.123
214.4104 |  | 211.124
215.0967 |  | 211.1241
210.8799 |  | 211.126
213.9727 |  | 211.126
213.96 |  | 211.1284
214.3149 |  | 211.1447
214.1615 |  | 211.1615
214.3625 |  | 211.1622
215.9646 |  | 211.17
213.6873 |  | 211.1715
212.9463 |  | 211.1737
212.8319 |  | 211.184
211.7219 |  | 211.185
212.2961 |  | 211.2004
214.2652 |  | 211.2013
214.237 |  | 211.2019
213.556 |  | 211.2037
213.8613 |  | 211.2093
214.1014 |  | 211.2148
213.1689 |  | 211.2192
213.5682 |  | 211.2206
213.7217 |  | 211.2237
214.3031 |  | 211.2285
213.67 |  | 211.2328
213.0829 |  | 211.2347
214.3672 |  | 211.2452
214.6823 |  | 211.2516
214.0056 |  | 211.2532
214.6037 |  | 211.2545
214.9737 |  | 211.257
215.0278 |  | 211.2615
211.7939 |  | 211.2648
214.043 |  | 211.2674
212.5754 |  | 211.2679
212.4471 |  | 211.2685
213.3477 |  | 211.2733
212.5562 |  | 211.2785
212.3315 |  | 211.2873
210.9977 |  | 211.2925
213.6197 |  | 211.2937
212.4758 |  | 211.2981
214.145 |  | 211.3018
214.2976 |  | 211.3084
211.7834 |  | 211.3084
216.0092 |  | 211.3102
215.1813 |  | 211.3111
214.2007 |  | 211.3129
214.088 |  | 211.3139
213.5571 |  | 211.3176
213.8739 |  | 211.3179
212.659 |  | 211.3179
215.2352 |  | 211.3233
213.0789 |  | 211.33
213.3423 |  | 211.3369
212.0305 |  | 211.3394
214.1789 |  | 211.3404
211.8912 |  | 211.3405
212.8281 |  | 211.3431
214.8811 |  | 211.3435
212.7983 |  | 211.3452
214.1537 |  | 211.3472
213.4579 |  | 211.3497
213.989 |  | 211.3516
214.6228 |  | 211.3519
212.3224 |  | 211.3523
212.1351 |  | 211.353
214.6944 |  | 211.3563
213.6586 |  | 211.3616
215.6802 |  | 211.3616
215.0357 |  | 211.3633
214.1382 |  | 211.3648
212.9992 |  | 211.3649
212.0995 |  | 211.365
214.7686 |  | 211.3683
213.2302 |  | 211.37
213.854 |  | 211.3747
212.6519 |  | 211.3774
213.4552 |  | 211.3794
211.2192 |  | 211.3801
213.7331 |  | 211.3803
212.7274 |  | 211.3821
214.0957 |  | 211.385
213.8931 |  | 211.3886
214.1558 |  | 211.3888
212.0697 |  | 211.3898
213.377 |  | 211.3899
212.6879 |  | 211.3933
213.2475 |  | 211.3934
213.8213 |  | 211.3976
212.8628 |  | 211.3991
212.8345 |  | 211.4003
212.2677 |  | 211.4034
213.1498 |  | 211.407
213.566 |  | 211.4105
212.4529 |  | 211.412
214.104 |  | 211.4132
213.1838 |  | 211.4139
213.7159 |  | 211.4146
213.5031 |  | 211.4151
213.0965 |  | 211.4185
214.1614 |  | 211.426
212.9126 |  | 211.4275
212.292 |  | 211.4331
214.8135 |  | 211.436
213.9992 |  | 211.4373
212.0556 |  | 211.4375
213.3312 |  | 211.4378
214.7801 |  | 211.4404
213.9825 |  | 211.4423
213.8808 |  | 211.4449
213.3737 |  | 211.4456
212.8303 |  | 211.4467
213.101 |  | 211.4503
211.8405 |  | 211.452
213.0458 |  | 211.4524
213.8557 |  | 211.4541
214.2552 |  | 211.4596
212.5285 |  | 211.4617
213.9811 |  | 211.4631
212.4936 |  | 211.4676
212.7171 |  | 211.4696
212.0428 |  | 211.4733
214.4921 |  | 211.4737
212.9096 |  | 211.4739
213.0496 |  | 211.4783
212.9606 |  | 211.4804
214.7701 |  | 211.4814
215.2459 |  | 211.4826
212.3432 |  | 211.4837
212.4463 |  | 211.488
215.3939 |  | 211.4881
213.0652 |  | 211.4911
213.0167 |  | 211.4972
211.096 |  | 211.4977
214.3223 |  | 211.4978
211.7049 |  | 211.4998
212.0064 |  | 211.5018
212.8421 |  | 211.5028
214.4898 |  | 211.5045
212.2616 |  | 211.5047
213.4622 |  | 211.5058
213.0811 |  | 211.5075
212.6235 |  | 211.5087
214.2072 |  | 211.5087
213.0375 |  | 211.5101
215.7288 |  | 211.5112
211.8456 |  | 211.5142
211.3404 |  | 211.5145
213.7359 |  | 211.5151
212.7907 |  | 211.5174
214.0618 |  | 211.5176
212.2001 |  | 211.5209
213.6653 |  | 211.5212
210.1315 |  | 211.525
212.7804 |  | 211.5273
212.0522 |  | 211.5313
214.2657 |  | 211.5345
213.1317 |  | 211.5375
212.4653 |  | 211.5375
213.4396 |  | 211.5401
212.6693 |  | 211.5437
213.1203 |  | 211.5444
214.4866 |  | 211.5476
212.4047 |  | 211.549
213.8841 |  | 211.5496
213.43 |  | 211.5499
213.2389 |  | 211.5501
211.4423 |  | 211.5502
215.5675 |  | 211.551
214.135 |  | 211.5558
213.2966 |  | 211.5588
213.6177 |  | 211.5642
213.7611 |  | 211.5674
214.4132 |  | 211.5675
213.5531 |  | 211.5694
213.6298 |  | 211.5711
212.4782 |  | 211.5724
212.428 |  | 211.5752
212.5867 |  | 211.5762
212.6643 |  | 211.5786
214.3991 |  | 211.5792
213.4728 |  | 211.5846
214.459 |  | 211.5846
215.4595 |  | 211.5897
214.1313 |  | 211.5931
212.5312 |  | 211.5937
213.5328 |  | 211.5985
215.4977 |  | 211.5989
211.9065 |  | 211.599
215.0995 |  | 211.6019
214.2173 |  | 211.6022
212.8658 |  | 211.6029
211.6943 |  | 211.604
212.1977 |  | 211.6044
213.8267 |  | 211.6063
214.0955 |  | 211.6103
215.7216 |  | 211.613
213.4306 |  | 211.6164
212.7798 |  | 211.6223
216.5399 |  | 211.6241
214.9802 |  | 211.6264
213.1125 |  | 211.6271
213.318 |  | 211.6297
213.71 |  | 211.6311
211.6903 |  | 211.6316
214.7079 |  | 211.6357
213.9557 |  | 211.636
214.1622 |  | 211.6367
215.3892 |  | 211.6414
213.1766 |  | 211.6442
211.8648 |  | 211.6463
213.8558 |  | 211.6489
214.2878 |  | 211.6492
214.2354 |  | 211.6501
212.6772 |  | 211.6501
214.1271 |  | 211.6538
214.3134 |  | 211.6572
213.2998 |  | 211.6577
211.3933 |  | 211.6596
213.6631 |  | 211.6607
212.8903 |  | 211.6616
212.1129 |  | 211.6633
214.5403 |  | 211.6637
214.4344 |  | 211.6647
213.8653 |  | 211.6653
213.3144 |  | 211.6664
214.1308 |  | 211.6687
213.6381 |  | 211.6688
213.7873 |  | 211.6689
211.385 |  | 211.6693
211.5087 |  | 211.6694
211.5176 |  | 211.6696
213.339 |  | 211.6699
211.488 |  | 211.6709
212.4935 |  | 211.6737
213.0878 |  | 211.6743
211.0786 |  | 211.6764
213.8183 |  | 211.6791
213.4704 |  | 211.6825
211.3111 |  | 211.6827
214.7763 |  | 211.6836
212.5315 |  | 211.6845
214.2059 |  | 211.6848
214.9315 |  | 211.6868
213.7384 |  | 211.6887
212.7636 |  | 211.6903
211.599 |  | 211.6911
213.8267 |  | 211.692
212.0232 |  | 211.6942
213.1209 |  | 211.6943
211.8753 |  | 211.697
214.9845 |  | 211.6974
214.7503 |  | 211.6974
212.4826 |  | 211.7018
213.0624 |  | 211.7031
212.0023 |  | 211.704
216.837 |  | 211.7041
213.3765 |  | 211.7049
214.6941 |  | 211.7073
211.4003 |  | 211.7124
211.4998 |  | 211.7141
213.5471 |  | 211.7159
212.6268 |  | 211.7161
213.6521 |  | 211.7176
214.4514 |  | 211.718
213.8912 |  | 211.7182
215.8912 |  | 211.7204
214.9321 |  | 211.721
212.2169 |  | 211.7214
213.5278 |  | 211.7217
212.1905 |  | 211.7219
212.2846 |  | 211.722
212.1301 |  | 211.7231
214.1349 |  | 211.7234
213.4429 |  | 211.7245
214.9539 |  | 211.726
214.3282 |  | 211.7281
212.7184 |  | 211.7309
212.4526 |  | 211.7314
214.7892 |  | 211.7321
212.5383 |  | 211.7323
212.7856 |  | 211.7351
213.2568 |  | 211.7356
212.7806 |  | 211.7367
213.9031 |  | 211.7432
213.057 |  | 211.7448
213.7369 |  | 211.7455
212.5258 |  | 211.746
214.1122 |  | 211.7473
214.4092 |  | 211.7477
214.8155 |  | 211.7525
214.0534 |  | 211.76
213.1331 |  | 211.7605
212.6609 |  | 211.7651
212.9427 |  | 211.7659
213.5492 |  | 211.7671
212.4474 |  | 211.7698
213.628 |  | 211.7708
213.4698 |  | 211.7717
214.3836 |  | 211.7728
212.6026 |  | 211.7734
212.6403 |  | 211.7734
213.6087 |  | 211.7745
212.9333 |  | 211.7745
214.255 |  | 211.7768
212.6551 |  | 211.7796
212.7565 |  | 211.7799
213.9556 |  | 211.7804
213.3241 |  | 211.7814
212.0192 |  | 211.7816
212.2512 |  | 211.7834
211.4331 |  | 211.7864
214.6664 |  | 211.7885
214.2206 |  | 211.7888
213.1481 |  | 211.789
213.2936 |  | 211.7896
213.1243 |  | 211.7902
213.01 |  | 211.7904
214.1488 |  | 211.7913
212.9077 |  | 211.7939
213.3467 |  | 211.794
212.2591 |  | 211.795
212.8927 |  | 211.7956
214.7059 |  | 211.7995
212.976 |  | 211.8006
213.7548 |  | 211.8012
212.2398 |  | 211.8015
215.6976 |  | 211.8027
213.5525 |  | 211.8033
212.5622 |  | 211.805
212.7563 |  | 211.8052
212.2196 |  | 211.8053
212.2881 |  | 211.8067
213.2695 |  | 211.8069
214.5242 |  | 211.8069
212.6315 |  | 211.808
212.7992 |  | 211.8093
212.5063 |  | 211.8103
212.2817 |  | 211.8117
212.8762 |  | 211.8148
213.3889 |  | 211.8161
212.6334 |  | 211.8182
213.091 |  | 211.8249
213.7778 |  | 211.8266
212.7712 |  | 211.827
213.0529 |  | 211.8274
212.9462 |  | 211.829
212.5094 |  | 211.83
212.5297 |  | 211.8301
212.7858 |  | 211.8313
212.2313 |  | 211.8329
214.441 |  | 211.8339
213.0585 |  | 211.8381
211.9114 |  | 211.8405
210.4392 |  | 211.8411
212.4258 |  | 211.8431
212.3832 |  | 211.8441
213.5816 |  | 211.8456
213.0332 |  | 211.8459
212.5908 |  | 211.8469
212.9697 |  | 211.8471
213.6329 |  | 211.8499
211.892 |  | 211.852
213.5117 |  | 211.852
213.2117 |  | 211.8533
214.8635 |  | 211.8551
213.3073 |  | 211.8553
213.7662 |  | 211.8559
212.6147 |  | 211.856
214.1728 |  | 211.8592
213.6076 |  | 211.8618
213.9041 |  | 211.8635
213.2563 |  | 211.8639
213.0722 |  | 211.8648
212.8011 |  | 211.8649
214.946 |  | 211.8651
211.9964 |  | 211.866
212.9872 |  | 211.8681
213.5277 |  | 211.8717
214.369 |  | 211.8725
212.3779 |  | 211.8729
213.8373 |  | 211.8732
214.6629 |  | 211.8747
214.2088 |  | 211.8753
212.8909 |  | 211.8768
213.3395 |  | 211.8773
213.6503 |  | 211.8794
213.406 |  | 211.8801
212.6913 |  | 211.8808
213.4849 |  | 211.8841
215.0976 |  | 211.8842
213.5225 |  | 211.8843
213.9533 |  | 211.8852
212.711 |  | 211.8858
214.8033 |  | 211.8876
212.5919 |  | 211.891
215.2481 |  | 211.8912
213.9031 |  | 211.8917
213.4071 |  | 211.892
213.5641 |  | 211.8947
212.1701 |  | 211.8953
214.0074 |  | 211.8969
213.5753 |  | 211.8975
214.1295 |  | 211.8989
213.6708 |  | 211.8998
211.3369 |  | 211.9024
211.4733 |  | 211.9027
212.8605 |  | 211.9031
214.8747 |  | 211.9034
213.7495 |  | 211.9065
212.2415 |  | 211.9075
214.3035 |  | 211.9087
214.1274 |  | 211.9114
213.8341 |  | 211.9115
214.0112 |  | 211.9122
212.5693 |  | 211.915
212.0214 |  | 211.9157
213.044 |  | 211.9181
212.3725 |  | 211.923
214.8596 |  | 211.9231
213.9634 |  | 211.9241
212.9911 |  | 211.9242
213.8387 |  | 211.9243
212.5713 |  | 211.9261
213.3155 |  | 211.9279
213.9077 |  | 211.9283
213.5465 |  | 211.9304
211.8381 |  | 211.9332
212.0323 |  | 211.9335
213.8732 |  | 211.935
209.7004 |  | 211.935
213.4226 |  | 211.9358
213.4707 |  | 211.9383
212.8825 |  | 211.9396
213.368 |  | 211.9449
214.2471 |  | 211.9473
211.9075 |  | 211.9482
213.688 |  | 211.9499
212.945 |  | 211.9501
213.7744 |  | 211.9503
213.5666 |  | 211.9539
212.8128 |  | 211.954
213.0741 |  | 211.9554
214.564 |  | 211.957
214.3381 |  | 211.9582
212.5067 |  | 211.961
212.5533 |  | 211.9625
213.6656 |  | 211.9632
213.4922 |  | 211.9638
211.8012 |  | 211.9638
212.7426 |  | 211.9649
212.921 |  | 211.9649
213.384 |  | 211.9655
211.7717 |  | 211.9694
212.3509 |  | 211.9699
213.4631 |  | 211.9706
213.633 |  | 211.9718
213.3134 |  | 211.9722
213.9659 |  | 211.9729
215.6575 |  | 211.9743
212.2689 |  | 211.9755
212.3315 |  | 211.9756
212.1253 |  | 211.9767
211.3683 |  | 211.9774
212.2985 |  | 211.9781
214.693 |  | 211.9782
214.791 |  | 211.9784
211.2937 |  | 211.9813
213.0352 |  | 211.9815
214.731 |  | 211.9815
215.0502 |  | 211.983
213.9016 |  | 211.9831
210.484 |  | 211.9843
212.9622 |  | 211.9843
214.1338 |  | 211.9844
211.8876 |  | 211.9849
215.4223 |  | 211.9856
213.4947 |  | 211.9872
211.7734 |  | 211.9911
216.141 |  | 211.9915
213.4948 |  | 211.9919
211.3899 |  | 211.9937
213.0051 |  | 211.9945
213.1046 |  | 211.9948
212.3069 |  | 211.9951
212.4011 |  | 211.9951
212.6201 |  | 211.9964
212.0088 |  | 211.9992
213.3485 |  | 212
211.3801 |  | 212.0008
214.3069 |  | 212.001
212.6856 |  | 212.0011
215.0152 |  | 212.0023
213.8707 |  | 212.0024
214.7998 |  | 212.0059
212.464 |  | 212.0064
214.0629 |  | 212.0071
215.3645 |  | 212.0088
212.8361 |  | 212.0097
212.3024 |  | 212.0111
213.3239 |  | 212.0133
212.3252 |  | 212.0135
212.6634 |  | 212.0138
214.2684 |  | 212.0141
213.2863 |  | 212.0146
212.2783 |  | 212.0146
211.124 |  | 212.0152
212.9325 |  | 212.0192
213.5045 |  | 212.0199
213.0795 |  | 212.0214
212.732 |  | 212.0232
211.4139 |  | 212.0256
213.0347 |  | 212.0287
212.345 |  | 212.0288
211.8969 |  | 212.0296
212.2151 |  | 212.0305
214.4307 |  | 212.0323
214.7388 |  | 212.0329
211.697 |  | 212.034
212.1247 |  | 212.0348
213.3878 |  | 212.0361
212.9525 |  | 212.0364
214.3366 |  | 212.0372
212.7428 |  | 212.0378
212.7655 |  | 212.0399
211.1221 |  | 212.0399
210.7991 |  | 212.0405
213.029 |  | 212.0413
214.413 |  | 212.0414
214.5429 |  | 212.0425
212.7544 |  | 212.0428
214.1653 |  | 212.043
213.6984 |  | 212.0441
213.3609 |  | 212.0456
213.6537 |  | 212.0457
213.9257 |  | 212.0462
211.4631 |  | 212.047
212.5742 |  | 212.0474
214.6492 |  | 212.0477
215.8174 |  | 212.0487
215.4397 |  | 212.049
212.3945 |  | 212.0504
212.1257 |  | 212.0522
213.9577 |  | 212.0531
211.9655 |  | 212.0533
214.2124 |  | 212.055
214.1383 |  | 212.0555
213.3358 |  | 212.0556
212.9619 |  | 212.0556
213.1476 |  | 212.0556
214.3987 |  | 212.0566
213.793 |  | 212.0574
212.2858 |  | 212.0575
213.3191 |  | 212.0577
213.8161 |  | 212.0619
211.704 |  | 212.0624
211.3497 |  | 212.0635
212.2648 |  | 212.0643
213.7581 |  | 212.065
214.6541 |  | 212.0665
212.5539 |  | 212.0678
214.1278 |  | 212.0694
213.8109 |  | 212.0696
213.0709 |  | 212.0697
212.0877 |  | 212.0713
212.1437 |  | 212.0718
214.0951 |  | 212.0721
214.1438 |  | 212.0734
212.9108 |  | 212.0761
213.3708 |  | 212.078
214.622 |  | 212.0789
211.064 |  | 212.0792
213.5082 |  | 212.0793
212.9263 |  | 212.0793
212.575 |  | 212.082
212.199 |  | 212.0834
215.2614 |  | 212.0836
213.84 |  | 212.0871
212.1609 |  | 212.0876
213.4411 |  | 212.0877
213.3363 |  | 212.088
212.9972 |  | 212.0882
212.7443 |  | 212.0883
214.7553 |  | 212.0883
213.2686 |  | 212.0888
211.9649 |  | 212.0895
213.7855 |  | 212.0921
210.6478 |  | 212.0922
214.7847 |  | 212.0932
213.8966 |  | 212.0933
213.0757 |  | 212.0935
212.718 |  | 212.0944
214.712 |  | 212.0949
212.6027 |  | 212.099
212.8111 |  | 212.0995
214.0516 |  | 212.101
212.266 |  | 212.1036
213.2215 |  | 212.1049
212.8231 |  | 212.1053
211.8635 |  | 212.1065
213.2609 |  | 212.1076
212.0504 |  | 212.108
212.8753 |  | 212.1083
214.9836 |  | 212.1087
211.4449 |  | 212.1088
214.7411 |  | 212.1091
213.2533 |  | 212.1095
212.9128 |  | 212.1095
213.9244 |  | 212.1101
212.2389 |  | 212.1107
214.9534 |  | 212.1119
213.4482 |  | 212.1129
213.3101 |  | 212.113
213.5353 |  | 212.1132
212.9065 |  | 212.1133
212.8007 |  | 212.1136
213.0978 |  | 212.1157
215.6448 |  | 212.1161
213.7844 |  | 212.1188
212.6433 |  | 212.1219
212.8013 |  | 212.1221
211.17 |  | 212.1223
213.3196 |  | 212.1226
212.5322 |  | 212.1245
213.8746 |  | 212.1247
214.0913 |  | 212.1253
212.1088 |  | 212.1257
211.7804 |  | 212.127
213.036 |  | 212.1273
212.9601 |  | 212.1301
212.3384 |  | 212.1311
212.9553 |  | 212.132
210.432 |  | 212.1333
213.7629 |  | 212.1345
214.7587 |  | 212.1351
211.9115 |  | 212.1361
212.2857 |  | 212.1375
211.7473 |  | 212.138
212.034 |  | 212.1384
214.5409 |  | 212.1387
214.3085 |  | 212.1394
211.7659 |  | 212.1423
214.6009 |  | 212.1429
212.675 |  | 212.1437
213.3935 |  | 212.1471
213.1166 |  | 212.1474
214.0241 |  | 212.1479
212.3488 |  | 212.1487
212.8395 |  | 212.1525
213.9502 |  | 212.1527
212.3632 |  | 212.1528
210.8393 |  | 212.1554
213.8834 |  | 212.1555
213.2528 |  | 212.1572
213.0189 |  | 212.1573
214.1829 |  | 212.158
211.9482 |  | 212.1584
213.516 |  | 212.1591
211.5711 |  | 212.1592
211.3821 |  | 212.1594
212.7761 |  | 212.1599
212.588 |  | 212.1599
214.4368 |  | 212.1602
212.7343 |  | 212.1609
211.6492 |  | 212.1617
213.6742 |  | 212.1623
212.27 |  | 212.1628
213.0123 |  | 212.1637
211.2452 |  | 212.1644
211.8067 |  | 212.1659
213.8606 |  | 212.166
211.795 |  | 212.1673
214.0087 |  | 212.1683
211.6103 |  | 212.1691
211.4783 |  | 212.1694
214.3037 |  | 212.1696
213.4442 |  | 212.1697
213.8029 |  | 212.1701
213.093 |  | 212.1707
214.4321 |  | 212.1709
214.0342 |  | 212.1724
212.8919 |  | 212.1731
215.2266 |  | 212.1735
215.7204 |  | 212.175
213.697 |  | 212.1751
214.571 |  | 212.1755
213.1734 |  | 212.1766
213.7761 |  | 212.1777
215.6617 |  | 212.178
212.5305 |  | 212.1783
212.8923 |  | 212.1794
213.2602 |  | 212.1806
213.3789 |  | 212.1811
214.1583 |  | 212.1824
213.4325 |  | 212.1847
214.1886 |  | 212.1848
213.4937 |  | 212.1887
213.1786 |  | 212.1894
212.4715 |  | 212.1895
212.7883 |  | 212.1905
214.0708 |  | 212.1912
212.2927 |  | 212.1923
213.2418 |  | 212.1954
213.6057 |  | 212.1977
212.284 |  | 212.1982
212.3391 |  | 212.199
214.3178 |  | 212.2001
213.1713 |  | 212.2022
214.0735 |  | 212.2046
212.5353 |  | 212.205
211.829 |  | 212.2054
214.2939 |  | 212.2059
212.6451 |  | 212.2061
213.3613 |  | 212.2061
213.9413 |  | 212.2072
213.8145 |  | 212.2075
213.1881 |  | 212.2077
214.3714 |  | 212.2081
214.7507 |  | 212.21
212.2255 |  | 212.2102
213.1097 |  | 212.2119
213.3819 |  | 212.2125
216.1368 |  | 212.2137
212.7275 |  | 212.2137
213.8544 |  | 212.2151
212.6268 |  | 212.2154
213.0613 |  | 212.2161
212.2137 |  | 212.2165
212.4547 |  | 212.2169
212.1223 |  | 212.218
214.1103 |  | 212.2196
214.9359 |  | 212.2218
211.7671 |  | 212.224
213.6351 |  | 212.2255
212.2852 |  | 212.2277
212.7867 |  | 212.2279
213.7329 |  | 212.2313
212.4619 |  | 212.2323
214.5654 |  | 212.2351
215.8188 |  | 212.2355
212.9838 |  | 212.2355
214.0849 |  | 212.2374
212.6348 |  | 212.2388
212.1423 |  | 212.2389
211.8329 |  | 212.2396
215.9218 |  | 212.2398
213.0981 |  | 212.2415
211.8551 |  | 212.2415
213.0947 |  | 212.2423
213.7075 |  | 212.2424
216.6535 |  | 212.2436
213.0902 |  | 212.2442
213.9451 |  | 212.2443
215.7529 |  | 212.2459
214.3703 |  | 212.2469
211.7314 |  | 212.2472
214.7444 |  | 212.2476
212.7656 |  | 212.2482
211.6596 |  | 212.2483
212.3321 |  | 212.2494
213.2347 |  | 212.2494
213.4868 |  | 212.2497
213.7638 |  | 212.2508
212.5149 |  | 212.2512
213.0181 |  | 212.2525
213.9916 |  | 212.2527
214.5576 |  | 212.253
214.3367 |  | 212.2566
212.1709 |  | 212.2567
213.5235 |  | 212.2582
213.1783 |  | 212.2584
212.2396 |  | 212.2587
211.3176 |  | 212.2587
212.9564 |  | 212.2591
212.5686 |  | 212.2597
213.7638 |  | 212.26
214.0818 |  | 212.2616
214.3978 |  | 212.2635
213.7191 |  | 212.2648
214.0652 |  | 212.2649
213.944 |  | 212.266
213.923 |  | 212.2677
214.1484 |  | 212.2689
212.386 |  | 212.27
213.2231 |  | 212.2717
214.5598 |  | 212.2741
214.2643 |  | 212.2764
211.9181 |  | 212.2767
214.18 |  | 212.2768
213.9691 |  | 212.2782
213.3537 |  | 212.2783
211.549 |  | 212.2794
212.3177 |  | 212.2817
214.7946 |  | 212.2829
212.5691 |  | 212.2836
212.4594 |  | 212.284
213.0161 |  | 212.2846
212.5823 |  | 212.2852
213.3445 |  | 212.2852
215.222 |  | 212.2857
214.6127 |  | 212.2858
211.8801 |  | 212.2881
214.2812 |  | 212.2895
213.5806 |  | 212.2903
214.0969 |  | 212.2907
212.9034 |  | 212.2915
212.3116 |  | 212.2915
212.4642 |  | 212.292
214.9157 |  | 212.2923
214.1821 |  | 212.2927
211.3794 |  | 212.296
214.6779 |  | 212.2961
212.2218 |  | 212.2985
212.6809 |  | 212.2996
213.7216 |  | 212.3021
213.2942 |  | 212.3024
213.1468 |  | 212.3033
212.4382 |  | 212.3038
214.6464 |  | 212.304
211.9499 |  | 212.3043
213.8986 |  | 212.3056
213.6633 |  | 212.3063
212.0793 |  | 212.3069
214.6142 |  | 212.3073
213.3904 |  | 212.3078
213.2537 |  | 212.3088
213.8361 |  | 212.3116
212.5319 |  | 212.3125
211.8773 |  | 212.3125
212.2494 |  | 212.3135
214.2705 |  | 212.3141
212.8572 |  | 212.3155
213.5915 |  | 212.3167
214.5222 |  | 212.3177
213.3725 |  | 212.3187
212.1751 |  | 212.3199
215.3278 |  | 212.3206
214.5167 |  | 212.3206
214.5468 |  | 212.321
213.0324 |  | 212.3217
214.1051 |  | 212.3219
209.8875 |  | 212.3224
212.7874 |  | 212.3235
214.2755 |  | 212.3237
214.004 |  | 212.3245
213.9078 |  | 212.3252
212.5269 |  | 212.3252
213.6147 |  | 212.3255
212.3951 |  | 212.3256
213.2891 |  | 212.3272
211.0795 |  | 212.3291
211.8947 |  | 212.3293
210.8205 |  | 212.3311
211.9774 |  | 212.3312
212.7486 |  | 212.3315
214.0217 |  | 212.3315
212.7071 |  | 212.3321
213.4114 |  | 212.3326
212.8556 |  | 212.3365
213.0752 |  | 212.3367
213.4525 |  | 212.3372
214.5962 |  | 212.3381
213.2892 |  | 212.3384
212.2794 |  | 212.3391
215.1428 |  | 212.3399
214.3216 |  | 212.3399
212.4709 |  | 212.3426
212.5323 |  | 212.3432
214.6716 |  | 212.3432
212.2494 |  | 212.3438
211.6223 |  | 212.3444
213.0343 |  | 212.3448
213.6055 |  | 212.345
212.6823 |  | 212.3471
214.8219 |  | 212.3483
214.5785 |  | 212.3488
211.5058 |  | 212.349
213.7468 |  | 212.3507
215.309 |  | 212.3509
215.3734 |  | 212.3514
212.444 |  | 212.3515
212.2154 |  | 212.3521
214.2902 |  | 212.353
213.1216 |  | 212.3532
213.6944 |  | 212.3567
214.0436 |  | 212.3575
213.916 |  | 212.358
213.7341 |  | 212.3585
214.6851 |  | 212.3593
213.9713 |  | 212.3597
214.4715 |  | 212.3604
212.001 |  | 212.3608
209.0713 |  | 212.3608
213.1444 |  | 212.3618
216.2553 |  | 212.3619
214.1632 |  | 212.3625
213.3559 |  | 212.3632
212.9041 |  | 212.365
212.6013 |  | 212.3653
214.9531 |  | 212.3675
214.5932 |  | 212.3682
211.9729 |  | 212.371
212.3043 |  | 212.3717
214.0059 |  | 212.3725
216.4233 |  | 212.3747
211.9813 |  | 212.3748
214.6766 |  | 212.3752
212.3608 |  | 212.3753
213.7994 |  | 212.3759
212.1528 |  | 212.377
211.8808 |  | 212.3771
213.7551 |  | 212.3779
212.9489 |  | 212.3801
215.2488 |  | 212.3805
212.4101 |  | 212.3823
215.7193 |  | 212.383
213.3571 |  | 212.3832
212.7536 |  | 212.3837
213.7943 |  | 212.3842
214.8243 |  | 212.3844
212.5343 |  | 212.3847
214.0913 |  | 212.3855
214.5019 |  | 212.3855
214.6398 |  | 212.3857
212.3753 |  | 212.386
212 |  | 212.3877
214.1194 |  | 212.3892
214.0452 |  | 212.3898
210.9825 |  | 212.39
212.9884 |  | 212.3918
213.1527 |  | 212.392
212.5868 |  | 212.3925
213.2152 |  | 212.3934
213.0029 |  | 212.3942
213.6147 |  | 212.3943
213.9493 |  | 212.3945
214.1825 |  | 212.3951
215.5381 |  | 212.397
212.9925 |  | 212.3998
212.6425 |  | 212.3998
213.37 |  | 212.4007
214.0764 |  | 212.4008
211.4275 |  | 212.4011
211.126 |  | 212.4027
212.9526 |  | 212.4031
212.8008 |  | 212.4037
210.9256 |  | 212.4041
213.1632 |  | 212.4043
214.3177 |  | 212.4047
213.4286 |  | 212.4066
215.0068 |  | 212.4072
211.7728 |  | 212.4074
212.1683 |  | 212.4074
212.5751 |  | 212.4081
213.1732 |  | 212.4088
213.5691 |  | 212.4101
213.8995 |  | 212.4108
214.0196 |  | 212.4123
212.3381 |  | 212.4125
214.192 |  | 212.4138
215.476 |  | 212.4145
213.975 |  | 212.4151
213.8778 |  | 212.4158
214.4972 |  | 212.416
212.5297 |  | 212.4173
214.2044 |  | 212.4178
213.9834 |  | 212.4202
214.6734 |  | 212.4216
214.3457 |  | 212.4217
213.515 |  | 212.4219
213.2795 |  | 212.4224
213.9478 |  | 212.4231
212.6922 |  | 212.4236
216.0209 |  | 212.424
214.3699 |  | 212.4249
211.6501 |  | 212.4258
213.2527 |  | 212.4258
212.3942 |  | 212.4265
211.005 |  | 212.4265
214.8043 |  | 212.4269
211.1094 |  | 212.428
213.4401 |  | 212.428
214.3379 |  | 212.4284
212.2717 |  | 212.4285
214.1896 |  | 212.4286
213.0516 |  | 212.4286
211.5558 |  | 212.4293
213.1065 |  | 212.4294
214.2086 |  | 212.4311
213.5399 |  | 212.4319
213.1493 |  | 212.4323
215.8787 |  | 212.4325
214.1562 |  | 212.4344
213.6373 |  | 212.4346
213.5536 |  | 212.4351
213.523 |  | 212.436
213.8517 |  | 212.4363
212.3245 |  | 212.4376
213.495 |  | 212.4382
212.602 |  | 212.4393
213.1682 |  | 212.4395
213.9533 |  | 212.4401
211.8274 |  | 212.4405
212.858 |  | 212.4424
214.9836 |  | 212.443
213.0882 |  | 212.444
213.4999 |  | 212.4445
213.7724 |  | 212.4455
213.696 |  | 212.4462
213.7839 |  | 212.4463
211.2328 |  | 212.4468
212.9859 |  | 212.447
213.6473 |  | 212.4471
213.5423 |  | 212.4472
213.8849 |  | 212.4474
214.6816 |  | 212.4474
212.1894 |  | 212.4482
212.4792 |  | 212.4496
213.002 |  | 212.4503
214.7389 |  | 212.4509
213.5151 |  | 212.451
215.4295 |  | 212.4525
214.4644 |  | 212.4526
212.0713 |  | 212.4529
212.5564 |  | 212.4547
214.4039 |  | 212.4564
214.2889 |  | 212.4572
212.4043 |  | 212.458
213.1484 |  | 212.4594
214.5346 |  | 212.4609
213.1859 |  | 212.4611
213.8581 |  | 212.4615
212.8126 |  | 212.4619
214.8631 |  | 212.4622
215.0903 |  | 212.4624
212.9183 |  | 212.464
212.9664 |  | 212.4642
212.6564 |  | 212.4653
212.5677 |  | 212.4663
214.7244 |  | 212.4677
213.8553 |  | 212.4694
212.4031 |  | 212.4696
213.2573 |  | 212.4699
214.0541 |  | 212.4701
211.5694 |  | 212.4707
212.7188 |  | 212.4709
212.9412 |  | 212.4715
214.3534 |  | 212.4715
212.472 |  | 212.472
214.2762 |  | 212.4735
212.0922 |  | 212.4758
213.4248 |  | 212.4759
212.1755 |  | 212.4761
213.408 |  | 212.4768
211.7913 |  | 212.4782
212.0071 |  | 212.4782
212.5898 |  | 212.4792
212.845 |  | 212.4809
214.407 |  | 212.4813
214.921 |  | 212.4817
213.5672 |  | 212.4823
212.4008 |  | 212.4826
213.024 |  | 212.4826
214.0006 |  | 212.4833
212.2081 |  | 212.4834
216.778 |  | 212.4838
216.4317 |  | 212.4843
212.7637 |  | 212.4847
213.3011 |  | 212.4848
213.0833 |  | 212.485
213.536 |  | 212.4853
213.6232 |  | 212.4858
211.2013 |  | 212.4894
210.6214 |  | 212.4896
212.436 |  | 212.4897
215.2081 |  | 212.4901
214.3831 |  | 212.4903
212.9474 |  | 212.4913
214.6328 |  | 212.4924
213.2233 |  | 212.4935
213.899 |  | 212.4936
214.8125 |  | 212.494
213.558 |  | 212.4965
215.518 |  | 212.4968
214.5981 |  | 212.4972
212.043 |  | 212.4985
213.0213 |  | 212.5007
212.6334 |  | 212.5013
212.7206 |  | 212.502
213.1626 |  | 212.5024
212.2829 |  | 212.5027
211.9781 |  | 212.5053
214.5577 |  | 212.5063
212.1107 |  | 212.5067
213.145 |  | 212.5074
212.1592 |  | 212.508
214.6267 |  | 212.508
213.5009 |  | 212.5094
212.2741 |  | 212.5107
211.3803 |  | 212.5139
212.9418 |  | 212.5149
211.3129 |  | 212.5155
213.4431 |  | 212.5157
212.6933 |  | 212.5172
214.3337 |  | 212.5177
211.2674 |  | 212.5179
212.1599 |  | 212.5202
214.1388 |  | 212.5206
214.583 |  | 212.5207
213.516 |  | 212.5218
213.3617 |  | 212.5219
214.7429 |  | 212.523
214.6819 |  | 212.5233
210.6897 |  | 212.5238
214.7377 |  | 212.5244
210.7592 |  | 212.5252
213.0487 |  | 212.5258
214.4879 |  | 212.526
214.1533 |  | 212.5262
212.7919 |  | 212.5269
214.2444 |  | 212.5272
211.915 |  | 212.5273
211.8033 |  | 212.5277
214.1682 |  | 212.5284
213.8148 |  | 212.5285
211.7176 |  | 212.5289
213.192 |  | 212.5297
213.1906 |  | 212.5297
213.2113 |  | 212.5297
213.7064 |  | 212.5305
214.4707 |  | 212.5305
213.0866 |  | 212.5306
212.6915 |  | 212.5312
213.8005 |  | 212.5315
212.3898 |  | 212.5319
212.7122 |  | 212.5322
211.3616 |  | 212.5323
213.9066 |  | 212.5327
213.1635 |  | 212.5329
213.8716 |  | 212.5336
211.1241 |  | 212.5341
213.9214 |  | 212.5343
212.101 |  | 212.5345
212.0619 |  | 212.5348
211.0402 |  | 212.5351
214.8798 |  | 212.5353
213.9407 |  | 212.5354
214.4011 |  | 212.5374
212.8558 |  | 212.5383
212.9486 |  | 212.5389
213.3209 |  | 212.539
213.2198 |  | 212.539
212.2525 |  | 212.5408
212.5155 |  | 212.545
214.8346 |  | 212.545
212.6591 |  | 212.546
217.1 |  | 212.5474
213.6031 |  | 212.5476
215.118 |  | 212.5483
213.8528 |  | 212.5483
212.7285 |  | 212.5504
214.636 |  | 212.5519
212.4677 |  | 212.5528
214.774 |  | 212.5533
212.9674 |  | 212.5539
211.9358 |  | 212.5544
212.8666 |  | 212.5554
213.0631 |  | 212.5562
214.4771 |  | 212.5564
211.0384 |  | 212.5565
213.9439 |  | 212.5579
213.7739 |  | 212.5582
214.0911 |  | 212.5593
212.4985 |  | 212.5612
212.365 |  | 212.5612
214.0307 |  | 212.5622
212.7856 |  | 212.5628
214.0377 |  | 212.5636
213.7425 |  | 212.5651
211.8431 |  | 212.566
211.935 |  | 212.5664
212.5476 |  | 212.5669
214.6474 |  | 212.567
213.0652 |  | 212.5677
213.1114 |  | 212.5681
212.7609 |  | 212.5684
211.9231 |  | 212.5686
213.1532 |  | 212.5691
213.5227 |  | 212.5693
211.8559 |  | 212.5694
212.9418 |  | 212.5695
215.265 |  | 212.5695
212.0624 |  | 212.57
212.2059 |  | 212.57
211.37 |  | 212.5708
215.3603 |  | 212.5711
214.752 |  | 212.5713
212.0575 |  | 212.5715
212.5735 |  | 212.5735
214.8956 |  | 212.574
213.7455 |  | 212.5742
211.5846 |  | 212.575
210.9885 |  | 212.5751
214.004 |  | 212.5754
211.8161 |  | 212.5754
212.6002 |  | 212.5769
215.1403 |  | 212.577
213.9157 |  | 212.577
213.4848 |  | 212.5784
211.4146 |  | 212.579
214.4243 |  | 212.5799
214.8077 |  | 212.5806
215.0028 |  | 212.581
213.7394 |  | 212.5823
214.2596 |  | 212.5865
210.7916 |  | 212.5866
214.0467 |  | 212.5867
214.8889 |  | 212.5868
213.5889 |  | 212.5868
211.8768 |  | 212.5871
213.2965 |  | 212.5877
212.7818 |  | 212.588
215.0685 |  | 212.5882
213.1698 |  | 212.5886
213.306 |  | 212.5891
212.2061 |  | 212.5898
213.4997 |  | 212.5907
214.6501 |  | 212.5908
212.5341 |  | 212.5914
212.5684 |  | 212.5919
215.9002 |  | 212.5921
213.6774 |  | 212.5924
215.2811 |  | 212.5947
213.488 |  | 212.595
213.9349 |  | 212.5958
212.1659 |  | 212.5958
210.8082 |  | 212.5967
214.7418 |  | 212.5992
213.2669 |  | 212.6002
213.0716 |  | 212.6004
214.824 |  | 212.6007
212.6131 |  | 212.6013
211.6664 |  | 212.6018
213.9395 |  | 212.602
213.3148 |  | 212.6026
213.4811 |  | 212.6027
213.5492 |  | 212.6028
212.0288 |  | 212.6062
213.5841 |  | 212.6082
213.2606 |  | 212.6111
212.4833 |  | 212.6114
210.4072 |  | 212.6128
211.9554 |  | 212.6131
214.3056 |  | 212.614
212.5389 |  | 212.6145
214.7835 |  | 212.6147
214.6891 |  | 212.6162
212.7674 |  | 212.6178
215.0819 |  | 212.6187
214.6183 |  | 212.6201
211.4881 |  | 212.6205
212.3326 |  | 212.6207
213.8982 |  | 212.6213
213.9413 |  | 212.6222
214.0311 |  | 212.6228
211.4814 |  | 212.6235
212.1036 |  | 212.6238
214.5344 |  | 212.6244
213.6555 |  | 212.6251
214.2926 |  | 212.6254
212.6793 |  | 212.6268
212.7525 |  | 212.6268
214.2328 |  | 212.6289
212.3448 |  | 212.6289
212.5297 |  | 212.6296
213.3105 |  | 212.6303
212.3855 |  | 212.631
211.6607 |  | 212.6311
214.9932 |  | 212.6315
211.5502 |  | 212.6317
215.1584 |  | 212.633
211.4503 |  | 212.6334
210.6857 |  | 212.6334
214.2476 |  | 212.6335
212.4395 |  | 212.6337
211.2648 |  | 212.6338
212.8493 |  | 212.6338
213.6278 |  | 212.6344
212.9582 |  | 212.6348
213.1736 |  | 212.6354
211.7245 |  | 212.6357
212.7437 |  | 212.6365
212.4037 |  | 212.6369
212.4622 |  | 212.6403
213.9853 |  | 212.6414
212.6722 |  | 212.6425
212.1572 |  | 212.6427
212.8456 |  | 212.6431
212.3593 |  | 212.6433
214.6524 |  | 212.6436
214.0927 |  | 212.6449
214.2863 |  | 212.6451
212.2388 |  | 212.6459
214.5621 |  | 212.6462
212.8379 |  | 212.647
213.4819 |  | 212.6473
212.5877 |  | 212.6478
211.185 |  | 212.6487
212.2472 |  | 212.649
213.1639 |  | 212.6499
212.775 |  | 212.6501
215.029 |  | 212.6502
213.8367 |  | 212.6512
212.4202 |  | 212.6519
212.1245 |  | 212.6526
212.9042 |  | 212.6527
214.9835 |  | 212.6531
212.5582 |  | 212.6536
212.5554 |  | 212.6551
212.7866 |  | 212.6563
213.4391 |  | 212.6564
214.1308 |  | 212.6588
211.8069 |  | 212.659
214.4409 |  | 212.6591
213.0151 |  | 212.6591
213.5763 |  | 212.6599
213.1067 |  | 212.6606
215.391 |  | 212.6609
213.0038 |  | 212.6616
212.6857 |  | 212.6628
214.7516 |  | 212.6634
213.8834 |  | 212.6639
212.2054 |  | 212.6643
215.0902 |  | 212.6646
213.7487 |  | 212.6646
213.4852 |  | 212.6657
214.2254 |  | 212.6658
212.6462 |  | 212.6676
212.0792 |  | 212.6684
214.0781 |  | 212.6691
214.5373 |  | 212.6693
212.706 |  | 212.6693
212.7831 |  | 212.6698
211.5045 |  | 212.6722
211.9383 |  | 212.6735
213.1552 |  | 212.6736
212.4853 |  | 212.6743
212.1555 |  | 212.6743
213.6193 |  | 212.6748
214.3631 |  | 212.675
215.123 |  | 212.6753
213.3614 |  | 212.6758
213.1864 |  | 212.6771
212.4817 |  | 212.6772
215.6064 |  | 212.6793
214.2874 |  | 212.6801
212.0364 |  | 212.6804
212.6254 |  | 212.6809
212.8242 |  | 212.6818
215.0681 |  | 212.682
213.342 |  | 212.682
214.4077 |  | 212.6823
213.5366 |  | 212.6823
213.2279 |  | 212.6826
214.2277 |  | 212.6838
214.5002 |  | 212.6839
212.5345 |  | 212.6856
214.0605 |  | 212.6857
212.5669 |  | 212.6858
212.9113 |  | 212.6867
212.9474 |  | 212.6868
213.0471 |  | 212.6874
213.2842 |  | 212.6875
214.0559 |  | 212.6879
213.5373 |  | 212.6902
214.6276 |  | 212.6904
214.3143 |  | 212.6908
212.8133 |  | 212.691
214.4946 |  | 212.6913
211.8411 |  | 212.6915
214.559 |  | 212.6918
211.935 |  | 212.6922
213.3687 |  | 212.6933
212.4572 |  | 212.6956
212.2903 |  | 212.6956
211.6887 |  | 212.6962
213.7588 |  | 212.6964
213.634 |  | 212.6974
214.768 |  | 212.6979
215.7025 |  | 212.6998
212.2061 |  | 212.7002
214.1874 |  | 212.7022
212.8259 |  | 212.7027
214.4847 |  | 212.7054
212.5664 |  | 212.706
215.2261 |  | 212.706
213.3009 |  | 212.7065
213.6425 |  | 212.7066
210.6731 |  | 212.7071
214.3821 |  | 212.71
212.5636 |  | 212.7101
214.716 |  | 212.7105
214.8122 |  | 212.711
214.8411 |  | 212.7114
214.7173 |  | 212.7115
212.8306 |  | 212.712
212.6365 |  | 212.712
214.1914 |  | 212.7122
214.5437 |  | 212.7133
213.0166 |  | 212.7134
214.5084 |  | 212.7143
213.6546 |  | 212.7171
214.7511 |  | 212.7171
213.5414 |  | 212.7172
213.877 |  | 212.718
213.0888 |  | 212.7184
213.9429 |  | 212.7188
214.0443 |  | 212.7198
213.5569 |  | 212.7206
214.0128 |  | 212.7212
215.1919 |  | 212.7216
214.2414 |  | 212.7221
211.6577 |  | 212.7229
212.6028 |  | 212.7233
211.551 |  | 212.724
214.3687 |  | 212.7246
211.8729 |  | 212.7251
210.6238 |  | 212.7274
214.7374 |  | 212.7275
214.1792 |  | 212.7276
211.5762 |  | 212.7278
213.0786 |  | 212.7285
213.7324 |  | 212.7289
213.4021 |  | 212.729
211.6693 |  | 212.7291
214.1136 |  | 212.7299
213.4428 |  | 212.7306
213.2889 |  | 212.731
212.4088 |  | 212.7316
212.6478 |  | 212.732
211.3435 |  | 212.7322
213.9355 |  | 212.7333
213.8274 |  | 212.7334
212.9087 |  | 212.7335
212.2907 |  | 212.7343
214.8093 |  | 212.7352
214.0486 |  | 212.7381
212.6639 |  | 212.7381
214.1158 |  | 212.7383
212.9313 |  | 212.7386
212.0949 |  | 212.74
213.3034 |  | 212.7426
214.1078 |  | 212.7428
214.0997 |  | 212.7435
212.3235 |  | 212.7436
214.2411 |  | 212.7437
214.2598 |  | 212.7443
214.2852 |  | 212.7457
211.1615 |  | 212.7464
211.808 |  | 212.7476
212.986 |  | 212.7486
214.2431 |  | 212.7499
211.5212 |  | 212.7501
213.0455 |  | 212.7504
213.4033 |  | 212.7504
215.1434 |  | 212.7525
212.2587 |  | 212.7525
212.4468 |  | 212.7536
212.6823 |  | 212.7544
214.7701 |  | 212.7557
212.3943 |  | 212.7563
213.0494 |  | 212.7565
211.4911 |  | 212.7565
213.0851 |  | 212.7573
213.0616 |  | 212.7578
213.471 |  | 212.7581
212.883 |  | 212.7592
214.8802 |  | 212.7593
214.775 |  | 212.7596
212.5483 |  | 212.7609
213.0252 |  | 212.7612
214.5572 |  | 212.7633
214.9758 |  | 212.7636
215.3648 |  | 212.7637
212.5273 |  | 212.764
213.08 |  | 212.7646
213.4787 |  | 212.7652
212.8707 |  | 212.7655
211.6911 |  | 212.7656
213.1595 |  | 212.7657
215.8044 |  | 212.7668
212.2476 |  | 212.7674
213.7747 |  | 212.7686
212.6691 |  | 212.7693
214.331 |  | 212.7711
213.4687 |  | 212.7712
214.2018 |  | 212.7715
214.3571 |  | 212.7719
211.9718 |  | 212.7727
214.1574 |  | 212.773
212.7879 |  | 212.774
213.4398 |  | 212.7741
211.7351 |  | 212.7746
211.6974 |  | 212.775
214.1047 |  | 212.7755
214.934 |  | 212.7761
214.2355 |  | 212.7763
213.1361 |  | 212.7769
212.3063 |  | 212.777
213.3748 |  | 212.777
213.2738 |  | 212.777
214.197 |  | 212.7774
215.3355 |  | 212.7794
214.0322 |  | 212.7798
214.0057 |  | 212.7804
214.3208 |  | 212.7806
214.5763 |  | 212.7818
212.0138 |  | 212.782
211.7799 |  | 212.7831
214.4438 |  | 212.7833
212.963 |  | 212.7834
211.4837 |  | 212.7835
213.424 |  | 212.7846
212.0111 |  | 212.7851
212.5157 |  | 212.7852
212.353 |  | 212.7855
212.8648 |  | 212.7856
213.9427 |  | 212.7856
212.0921 |  | 212.7858
212.4108 |  | 212.7863
212.4258 |  | 212.7866
213.6431 |  | 212.7867
214.0683 |  | 212.7867
214.4651 |  | 212.7874
213.6239 |  | 212.7876
212.0024 |  | 212.7877
213.476 |  | 212.7879
214.2059 |  | 212.7883
214.4482 |  | 212.7887
212.0487 |  | 212.7888
213.0415 |  | 212.7907
215.5942 |  | 212.7912
214.9281 |  | 212.7919
212.447 |  | 212.7919
213.3391 |  | 212.7938
213.9853 |  | 212.7938
214.5969 |  | 212.7943
212.6628 |  | 212.7946
212.5238 |  | 212.7951
212.6449 |  | 212.797
214.3501 |  | 212.7975
212.4363 |  | 212.7981
213.2832 |  | 212.7983
212.4696 |  | 212.7983
213.4475 |  | 212.7992
210.9924 |  | 212.8003
213.7628 |  | 212.8005
212.9684 |  | 212.8007
213.1878 |  | 212.8008
212.8773 |  | 212.8011
213.1456 |  | 212.8013
214.7979 |  | 212.8018
213.4098 |  | 212.8023
214.4195 |  | 212.8026
213.9202 |  | 212.8031
214.2518 |  | 212.8033
211.8852 |  | 212.8041
213.3758 |  | 212.805
214.193 |  | 212.8057
212.8397 |  | 212.8061
213.0491 |  | 212.8065
214.4649 |  | 212.8088
213.9724 |  | 212.8101
212.4401 |  | 212.8102
214.4903 |  | 212.8104
211.8093 |  | 212.8108
211.8732 |  | 212.8109
210.6057 |  | 212.8111
213.9939 |  | 212.8115
213.1733 |  | 212.8118
214.6685 |  | 212.8126
211.6271 |  | 212.8128
213.7443 |  | 212.8133
213.2305 |  | 212.8143
214.4569 |  | 212.8185
213.8984 |  | 212.8199
214.4571 |  | 212.8199
211.9649 |  | 212.8209
211.5989 |  | 212.8209
213.7917 |  | 212.8231
213.2415 |  | 212.8238
213.3967 |  | 212.8239
211.9767 |  | 212.8242
214.5573 |  | 212.8242
213.079 |  | 212.8252
212.901 |  | 212.8259
212.8115 |  | 212.8271
212.6526 |  | 212.8273
211.8618 |  | 212.8281
211.6311 |  | 212.8284
212.1735 |  | 212.8286
214.155 |  | 212.8298
212.1474 |  | 212.8303
214.019 |  | 212.8306
211.4151 |  | 212.8316
213.3953 |  | 212.8319
213.4394 |  | 212.8342
213.2264 |  | 212.8345
213.3322 |  | 212.8347
214.9568 |  | 212.8361
212.9427 |  | 212.8361
212.9311 |  | 212.8366
212.5799 |  | 212.8371
213.6308 |  | 212.8379
213.5168 |  | 212.8395
214.4188 |  | 212.8397
214.0346 |  | 212.8421
211.5985 |  | 212.845
213.1707 |  | 212.8456
214.6249 |  | 212.846
213.9895 |  | 212.8461
213.8771 |  | 212.8466
211.9843 |  | 212.847
213.2951 |  | 212.8483
214.0161 |  | 212.8484
213.1662 |  | 212.8493
212.8065 |  | 212.8495
214.744 |  | 212.8505
214.828 |  | 212.8511
212.6162 |  | 212.8512
212.4405 |  | 212.8527
213.5399 |  | 212.8534
214.522 |  | 212.8537
213.2516 |  | 212.8539
212.4285 |  | 212.854
212.9451 |  | 212.8556
211.954 |  | 212.8558
215.0017 |  | 212.8564
213.3107 |  | 212.8572
214.4335 |  | 212.858
213.2943 |  | 212.8582
213.1398 |  | 212.8584
214.0898 |  | 212.8586
213.164 |  | 212.8587
213.9987 |  | 212.8601
211.525 |  | 212.8605
212.7133 |  | 212.8613
213.1219 |  | 212.8618
212.4231 |  | 212.8619
213.2419 |  | 212.8623
212.4609 |  | 212.8628
214.71 |  | 212.8636
212.1375 |  | 212.8648
212.3088 |  | 212.8653
213.216 |  | 212.8658
213.0852 |  | 212.8659
212.567 |  | 212.8663
214.6374 |  | 212.8666
212.5866 |  | 212.8666
215.3324 |  | 212.8667
213.3482 |  | 212.8669
211.1737 |  | 212.8679
212.5947 |  | 212.8682
213.2965 |  | 212.8682
214.1704 |  | 212.8689
214.466 |  | 212.8698
212.4072 |  | 212.8701
213.3435 |  | 212.8703
215.0342 |  | 212.8707
214.6839 |  | 212.8729
215.7002 |  | 212.8736
211.7234 |  | 212.8737
214.0446 |  | 212.8748
213.6753 |  | 212.8753
213.5507 |  | 212.8762
212.8209 |  | 212.877
214.7517 |  | 212.8771
212.428 |  | 212.8773
212.9731 |  | 212.8784
212.7105 |  | 212.8812
213.2126 |  | 212.8825
213.4152 |  | 212.8827
212.0574 |  | 212.883
212.7693 |  | 212.8844
214.3106 |  | 212.8859
215.148 |  | 212.8861
214.1954 |  | 212.8862
214.1801 |  | 212.8872
212.21 |  | 212.8883
215.0893 |  | 212.8901
211.8459 |  | 212.8903
213.1103 |  | 212.8906
213.5911 |  | 212.8909
214.4383 |  | 212.8912
211.2285 |  | 212.8919
212.0152 |  | 212.8923
213.1071 |  | 212.8927
212.6289 |  | 212.8927
211.4972 |  | 212.8929
211.257 |  | 212.8935
215.142 |  | 212.8937
212.0789 |  | 212.8941
213.713 |  | 212.8947
212.2469 |  | 212.8963
214.0762 |  | 212.897
211.9473 |  | 212.8971
212.8636 |  | 212.8976
212.9088 |  | 212.8979
213.1742 |  | 212.8991
213.5804 |  | 212.8995
213.786 |  | 212.901
212.6427 |  | 212.9015
212.8872 |  | 212.9029
212.1895 |  | 212.9034
214.1271 |  | 212.9041
213.6475 |  | 212.9042
212.108 |  | 212.9042
214.7183 |  | 212.9047
212.6693 |  | 212.9053
213.0426 |  | 212.9063
215.5504 |  | 212.9065
215.4276 |  | 212.9071
213.4144 |  | 212.9077
212.6588 |  | 212.9087
214.1047 |  | 212.9088
212.3167 |  | 212.9088
210.8363 |  | 212.9095
211.9815 |  | 212.9096
211.7448 |  | 212.9098
211.3991 |  | 212.9108
214.3999 |  | 212.9113
214.0051 |  | 212.9126
209.7003 |  | 212.9128
213.1649 |  | 212.9136
213.3711 |  | 212.9137
214.5597 |  | 212.9138
214.3956 |  | 212.914
214.0734 |  | 212.9154
213.4527 |  | 212.9157
212.9138 |  | 212.9159
215.4163 |  | 212.9162
212.458 |  | 212.9169
216.9912 |  | 212.9179
214.2178 |  | 212.9183
213.4594 |  | 212.9183
213.6861 |  | 212.9194
213.6342 |  | 212.9194
213.1977 |  | 212.9194
212.9029 |  | 212.9198
211.3888 |  | 212.921
211.5642 |  | 212.9217
215.427 |  | 212.923
211.7864 |  | 212.9245
212.1912 |  | 212.9254
213.3083 |  | 212.9257
213.1284 |  | 212.926
214.8484 |  | 212.9263
214.5425 |  | 212.9265
214.8358 |  | 212.9268
214.269 |  | 212.9278
213.864 |  | 212.9285
214.2669 |  | 212.9302
212.57 |  | 212.9306
211.365 |  | 212.9311
213.9658 |  | 212.9312
212.2022 |  | 212.9313
213.0759 |  | 212.9325
211.3179 |  | 212.9327
213.3361 |  | 212.9333
213.6333 |  | 212.9356
213.9737 |  | 212.9397
213.7931 |  | 212.9398
212.0135 |  | 212.9411
212.773 |  | 212.9412
213.1238 |  | 212.9418
211.5897 |  | 212.9418
214.6207 |  | 212.9427
212.7887 |  | 212.9427
212.8883 |  | 212.9433
211.636 |  | 212.9447
214.2066 |  | 212.945
212.3293 |  | 212.9451
214.7657 |  | 212.9456
211.3394 |  | 212.9462
213.786 |  | 212.9463
213.6939 |  | 212.947
213.3121 |  | 212.9474
212.3432 |  | 212.9474
211.6827 |  | 212.9474
212.947 |  | 212.9475
214.1076 |  | 212.9486
212.8619 |  | 212.9487
213.4085 |  | 212.9489
212.3925 |  | 212.9496
213.1101 |  | 212.9497
215.112 |  | 212.9501
209.5057 |  | 212.9514
213.1777 |  | 212.9517
215.2203 |  | 212.9524
214.7064 |  | 212.9525
213.672 |  | 212.9526
213.684 |  | 212.9528
213.4334 |  | 212.9532
213.4883 |  | 212.954
213.6392 |  | 212.9552
214.0894 |  | 212.9553
210.5025 |  | 212.9556
214.4831 |  | 212.9556
214.8851 |  | 212.9561
213.8828 |  | 212.9564
214.7005 |  | 212.9565
213.7382 |  | 212.9582
214.219 |  | 212.9592
214.9288 |  | 212.9601
213.826 |  | 212.9602
213.1831 |  | 212.9606
212.2895 |  | 212.9619
213.207 |  | 212.9622
215.0014 |  | 212.963
212.1811 |  | 212.9638
213.2603 |  | 212.9639
215.1555 |  | 212.9639
214.1895 |  | 212.9655
213.4395 |  | 212.9664
213.3265 |  | 212.9671
214.2881 |  | 212.9674
211.5724 |  | 212.9682
212.088 |  | 212.9684
211.5375 |  | 212.9686
212.9398 |  | 212.9697
213.1816 |  | 212.9702
212.1083 |  | 212.9703
212.5754 |  | 212.9703
212.6228 |  | 212.9711
214.8617 |  | 212.9729
212.8023 |  | 212.9731
213.096 |  | 212.9743
214.825 |  | 212.9748
214.4041 |  | 212.9758
214.4912 |  | 212.976
214.5821 |  | 212.9763
212.4834 |  | 212.9764
212.7436 |  | 212.9768
212.4027 |  | 212.9773
213.2586 |  | 212.9778
213.6329 |  | 212.9786
213.7404 |  | 212.9804
211.6316 |  | 212.9805
211.7708 |  | 212.9811
213.9707 |  | 212.9838
214.0232 |  | 212.9851
214.2217 |  | 212.9852
212.6311 |  | 212.9856
211.8989 |  | 212.9859
215.3489 |  | 212.986
212.5695 |  | 212.9866
211.0083 |  | 212.9872
212.383 |  | 212.9877
212.3155 |  | 212.9883
211.6019 |  | 212.9884
213.3315 |  | 212.989
214.6833 |  | 212.9911
211.5018 |  | 212.9925
214.4264 |  | 212.9931
211.7745 |  | 212.9936
212.8784 |  | 212.9944
213.7876 |  | 212.9948
212.877 |  | 212.9951
213.3275 |  | 212.9952
211.9638 |  | 212.996
214.126 |  | 212.9968
214.4576 |  | 212.997
212.6251 |  | 212.9972
212.9936 |  | 212.9979
213.7721 |  | 212.9986
213.7827 |  | 212.9992
213.2852 |  | 212.9992
211.8015 |  | 212.9997
212.8041 |  | 212.9997
213.4087 |  | 212.9998
214.121 |  | 212.9999
215.4552 |  | 213.0001
212.7278 |  | 213.0002
214.1245 |  | 213.0008
213.0673 |  | 213.0014
212.1525 |  | 213.002
211.5028 |  | 213.0029
212.4125 |  | 213.0035
212.7499 |  | 213.0038
212.0405 |  | 213.0043
213.155 |  | 213.0045
214.3604 |  | 213.0051
213.2064 |  | 213.0056
213.1719 |  | 213.01
212.4351 |  | 213.0108
213.7002 |  | 213.0123
211.8681 |  | 213.0133
211.8592 |  | 213.0136
213.0678 |  | 213.0138
211.76 |  | 213.0141
214.0227 |  | 213.0145
212.0477 |  | 213.0146
214.1988 |  | 213.0149
213.0462 |  | 213.015
213.3613 |  | 213.0151
212.8298 |  | 213.0161
213.2699 |  | 213.0166
213.804 |  | 213.0167
212.3125 |  | 213.0172
213.6422 |  | 213.0176
212.4894 |  | 213.0176
214.9642 |  | 213.0181
213.7451 |  | 213.0181
214.1415 |  | 213.0188
211.4696 |  | 213.0189
214.7589 |  | 213.019
212.6007 |  | 213.0191
211.6974 |  | 213.0195
212.0329 |  | 213.0213
213.701 |  | 213.0213
212.1384 |  | 213.0215
215.7818 |  | 213.024
213.4452 |  | 213.0252
213.1392 |  | 213.0256
215.8778 |  | 213.0266
215.4477 |  | 213.0284
213.1384 |  | 213.0287
213.4066 |  | 213.029
211.3233 |  | 213.0292
213.4421 |  | 213.0294
216.0576 |  | 213.0297
213.7097 |  | 213.0314
213.2074 |  | 213.0315
214.5092 |  | 213.032
212.7835 |  | 213.0322
213.3086 |  | 213.0324
212.3438 |  | 213.033
213.0652 |  | 213.0331
213.1789 |  | 213.0332
213.0292 |  | 213.0343
213.1017 |  | 213.0347
213.1246 |  | 213.0348
213.8916 |  | 213.0352
214.1215 |  | 213.036
213.3515 |  | 213.0362
213.8477 |  | 213.0366
212.4323 |  | 213.0369
211.7321 |  | 213.0375
212.2459 |  | 213.0375
212.7381 |  | 213.039
212.566 |  | 213.0394
214.7896 |  | 213.0396
213.3372 |  | 213.0404
213.9346 |  | 213.0404
212.6563 |  | 213.0415
212.3823 |  | 213.0425
212.4761 |  | 213.0426
211.746 |  | 213.0427
213.0176 |  | 213.044
215.8012 |  | 213.044
213.9672 |  | 213.0455
212.7711 |  | 213.0458
215.0001 |  | 213.0459
213.6013 |  | 213.0462
213.5796 |  | 213.0471
213.4999 |  | 213.0479
213.8085 |  | 213.0483
212.3918 |  | 213.0487
212.3855 |  | 213.0487
212.9856 |  | 213.0491
212.6317 |  | 213.0494
214.483 |  | 213.0494
212.6818 |  | 213.0496
216.342 |  | 213.0496
214.1719 |  | 213.05
212.9063 |  | 213.0516
213.3983 |  | 213.0518
212.3399 |  | 213.0519
214.6899 |  | 213.0522
213.0565 |  | 213.0529
212.9592 |  | 213.0565
212.2072 |  | 213.057
214.5882 |  | 213.0571
212.7352 |  | 213.0585
214.8253 |  | 213.0594
214.4848 |  | 213.06
213.9046 |  | 213.0604
214.3425 |  | 213.0611
213.2849 |  | 213.0612
213.1305 |  | 213.0613
213.5524 |  | 213.0616
213.381 |  | 213.0617
214.7457 |  | 213.0624
214.1254 |  | 213.0631
211.613 |  | 213.0652
215.0296 |  | 213.0652
214.2336 |  | 213.0652
214.0923 |  | 213.0673
213.8112 |  | 213.0678
214.5893 |  | 213.0687
212.4293 |  | 213.0687
213.2274 |  | 213.0692
212.579 |  | 213.0694
213.7347 |  | 213.0694
211.726 |  | 213.0699
213.0002 |  | 213.0708
214.0738 |  | 213.0709
213.3373 |  | 213.0713
213.0692 |  | 213.0716
212.0372 |  | 213.0719
212.5695 |  | 213.0722
213.9603 |  | 213.0738
213.1988 |  | 213.0741
212.4972 |  | 213.0744
215.7949 |  | 213.0748
214.1796 |  | 213.0752
213.3293 |  | 213.0755
213.9581 |  | 213.0757
215.3788 |  | 213.0759
212.1887 |  | 213.0786
213.5643 |  | 213.0786
212.7476 |  | 213.0789
214.1459 |  | 213.079
213.0314 |  | 213.0794
211.9911 |  | 213.0795
212.914 |  | 213.08
212.5784 |  | 213.0805
214.6133 |  | 213.0808
212.7646 |  | 213.081
213.0612 |  | 213.0811
213.9124 |  | 213.0817
214.4549 |  | 213.0818
212.6111 |  | 213.0829
214.6907 |  | 213.0831
211.9027 |  | 213.0832
213.7642 |  | 213.0833
215.8914 |  | 213.0851
212.3471 |  | 213.0852
210.8522 |  | 213.0858
212.1311 |  | 213.0861
211.6357 |  | 213.0866
215.155 |  | 213.0868
214.0696 |  | 213.0869
212.5528 |  | 213.0869
212.4081 |  | 213.0871
215.6443 |  | 213.0878
213.3391 |  | 213.0882
212.1087 |  | 213.0888
212.2374 |  | 213.0889
212.7876 |  | 213.0891
214.3801 |  | 213.0902
213.6725 |  | 213.091
213.7201 |  | 213.0914
212.3608 |  | 213.0914
212.4965 |  | 213.0919
213.5796 |  | 213.0928
211.6694 |  | 213.093
213.1958 |  | 213.0943
212.8937 |  | 213.0947
213.1927 |  | 213.096
212.8844 |  | 213.0963
213.5165 |  | 213.0965
212.9162 |  | 213.0967
214.8545 |  | 213.0978
213.8258 |  | 213.0981
211.4617 |  | 213.0984
212.6337 |  | 213.0991
212.8613 |  | 213.0997
213.9275 |  | 213.1008
210.5683 |  | 213.101
211.9304 |  | 213.101
212.4968 |  | 213.1017
213.2321 |  | 213.1021
213.4749 |  | 213.1046
214.3624 |  | 213.1056
211.7734 |  | 213.1065
214.2511 |  | 213.1067
212.3653 |  | 213.1069
213.2941 |  | 213.1071
212.6735 |  | 213.1097
213.9116 |  | 213.1101
213.2827 |  | 213.1103
212.7171 |  | 213.1114
212.3585 |  | 213.1116
213.0195 |  | 213.1121
210.6366 |  | 213.1125
211.7204 |  | 213.1146
213.8011 |  | 213.115
212.4847 |  | 213.1153
213.515 |  | 213.1155
213.8962 |  | 213.116
212.4319 |  | 213.1162
213.6652 |  | 213.1166
213.8925 |  | 213.1194
211.7651 |  | 213.1197
213.1547 |  | 213.1203
214.8879 |  | 213.1209
213.9637 |  | 213.1216
214.0398 |  | 213.1219
214.9149 |  | 213.1228
213.2914 |  | 213.123
212.6512 |  | 213.1238
213.1146 |  | 213.1243
213.4019 |  | 213.1244
212.996 |  | 213.1246
213.9736 |  | 213.1262
214.8561 |  | 213.1266
212.8199 |  | 213.1275
214.2486 |  | 213.1282
212.8537 |  | 213.1284
213.3312 |  | 213.1287
213.3624 |  | 213.1291
212.5806 |  | 213.1297
213.0294 |  | 213.1305
215.343 |  | 213.1317
212.7851 |  | 213.1322
212.7592 |  | 213.1322
212.8539 |  | 213.133
214.75 |  | 213.1331
214.3496 |  | 213.1335
213.5058 |  | 213.1337
211.7018 |  | 213.1338
213.0914 |  | 213.1345
214.1279 |  | 213.1359
214.6977 |  | 213.1361
213.8927 |  | 213.1365
212.0678 |  | 213.138
212.7719 |  | 213.1384
212.8748 |  | 213.1392
214.6024 |  | 213.1394
213.0331 |  | 213.1398
212.8859 |  | 213.1416
213.8418 |  | 213.1435
216.3815 |  | 213.144
213.8654 |  | 213.1444
213.5515 |  | 213.1445
213.5136 |  | 213.1446
213.3968 |  | 213.1447
213.2518 |  | 213.145
213.0869 |  | 213.1456
215.0766 |  | 213.1456
211.4977 |  | 213.1466
213.3259 |  | 213.1468
213.6463 |  | 213.1474
214.6832 |  | 213.1475
212.8033 |  | 213.1476
214.1998 |  | 213.1481
214.0172 |  | 213.1484
212.1091 |  | 213.1493
212.4066 |  | 213.1494
212.8108 |  | 213.1496
214.1777 |  | 213.1498
213.0861 |  | 213.1506
214.5238 |  | 213.152
212.9456 |  | 213.1527
214.823 |  | 213.1527
215.1601 |  | 213.153
215.3036 |  | 213.1532
213.3677 |  | 213.1537
213.3718 |  | 213.1547
212.8505 |  | 213.155
213.0738 |  | 213.1552
211.7768 |  | 213.156
212.7022 |  | 213.1564
213.2806 |  | 213.1569
213.1262 |  | 213.1578
215.6845 |  | 213.158
212.6414 |  | 213.158
213.0176 |  | 213.1595
214.0839 |  | 213.1596
213.8419 |  | 213.1599
212.5027 |  | 213.161
213.0928 |  | 213.1614
213.1466 |  | 213.1621
212.8663 |  | 213.1626
212.371 |  | 213.1631
215.382 |  | 213.1632
212.0556 |  | 213.1635
213.9833 |  | 213.1639
214.4643 |  | 213.164
213.8225 |  | 213.1641
211.6764 |  | 213.1646
212.7525 |  | 213.1649
215.1555 |  | 213.1662
212.502 |  | 213.167
213.3068 |  | 213.1677
212.5306 |  | 213.1682
213.9137 |  | 213.1689
213.2559 |  | 213.169
214.6862 |  | 213.1697
213.625 |  | 213.1698
212.1573 |  | 213.1699
213.4695 |  | 213.17
213.3169 |  | 213.1707
213.2468 |  | 213.1712
212.9763 |  | 213.1713
214.7555 |  | 213.1719
213.4113 |  | 213.1719
212.6128 |  | 213.1722
213.4358 |  | 213.1732
213.156 |  | 213.1733
214.448 |  | 213.1734
213.7949 |  | 213.1736
212.26 |  | 213.1742
211.8499 |  | 213.1759
211.2981 |  | 213.1766
215.3617 |  | 213.1768
213.7076 |  | 213.1777
214.2076 |  | 213.1779
212.8371 |  | 213.1783
213.1931 |  | 213.1784
213.8386 |  | 213.1786
213.3433 |  | 213.1789
214.5313 |  | 213.1815
211.9856 |  | 213.1816
213.2991 |  | 213.1821
214.2475 |  | 213.1822
215.3729 |  | 213.1825
214.0159 |  | 213.1825
213.1337 |  | 213.1831
212.6114 |  | 213.1838
213.2883 |  | 213.1843
213.0858 |  | 213.1845
212.3877 |  | 213.1859
213.4923 |  | 213.1864
211.4826 |  | 213.1869
212.4809 |  | 213.1878
214.1964 |  | 213.1879
214.7304 |  | 213.188
213.4398 |  | 213.1881
215.1426 |  | 213.1887
212.4455 |  | 213.1896
214.9452 |  | 213.1905
213.7571 |  | 213.1906
212.8995 |  | 213.192
212.8682 |  | 213.1927
212.0059 |  | 213.1931
215.5221 |  | 213.1945
211.6414 |  | 213.1952
214.5339 |  | 213.1958
214.9707 |  | 213.1977
213.1569 |  | 213.1985
212.1599 |  | 213.1988
211.8841 |  | 213.1997
213.5353 |  | 213.1998
213.1614 |  | 213.2003
211.5792 |  | 213.2005
211.5174 |  | 213.2012
214.945 |  | 213.2018
212.1527 |  | 213.2024
212.545 |  | 213.2033
214.2003 |  | 213.2054
213.8032 |  | 213.2054
212.5992 |  | 213.2064
212.539 |  | 213.2066
214.5738 |  | 213.2068
214.4984 |  | 213.207
213.6274 |  | 213.207
211.3139 |  | 213.2074
212.9997 |  | 213.2082
211.7031 |  | 213.2095
213.6132 |  | 213.2111
212.8102 |  | 213.2113
213.2976 |  | 213.2117
212.9952 |  | 213.212
212.731 |  | 213.2122
213.3094 |  | 213.2126
212.5474 |  | 213.2132
212.3219 |  | 213.2133
211.7896 |  | 213.2146
214.1167 |  | 213.2152
214.0394 |  | 213.216
214.5698 |  | 213.2173
213.4424 |  | 213.2183
214.3525 |  | 213.2196
214.2705 |  | 213.2198
215.2369 |  | 213.2211
213.4644 |  | 213.2215
212.6979 |  | 213.222
213.2886 |  | 213.2225
212.1623 |  | 213.2229
213.1359 |  | 213.2231
212.8703 |  | 213.2233
213.7513 |  | 213.2264
214.4058 |  | 213.2269
214.9148 |  | 213.2274
215.08 |  | 213.2274
213.5918 |  | 213.2276
214.0299 |  | 213.2279
213.1564 |  | 213.2282
212.7334 |  | 213.229
214.2063 |  | 213.2302
212.0199 |  | 213.2305
212.4707 |  | 213.2321
213.0871 |  | 213.2337
212.9851 |  | 213.2337
214.0663 |  | 213.2346
212.1049 |  | 213.2347
216.7765 |  | 213.2359
212.1188 |  | 213.2365
213.1153 |  | 213.2366
213.9645 |  | 213.237
215.2558 |  | 213.2373
212.7888 |  | 213.2388
214.7736 |  | 213.2389
214.535 |  | 213.2395
214.2433 |  | 213.2396
213.2404 |  | 213.2399
212.8061 |  | 213.2404
213.6639 |  | 213.2407
214.4073 |  | 213.2411
213.1677 |  | 213.2413
211.5937 |  | 213.2415
213.0138 |  | 213.2418
212.1594 |  | 213.2419
211.4456 |  | 213.2422
212.3073 |  | 213.2425
212.9053 |  | 213.2438
213.0496 |  | 213.2446
211.794 |  | 213.2447
214.1277 |  | 213.2468
212.5628 |  | 213.2472
215.5009 |  | 213.2475
213.2337 |  | 213.2498
214.9328 |  | 213.25
213.6484 |  | 213.2507
212.5694 |  | 213.251
211.2685 |  | 213.2513
214.27 |  | 213.2516
214.4676 |  | 213.2518
214.3507 |  | 213.2519
213.9057 |  | 213.2525
213.422 |  | 213.2527
213.7705 |  | 213.2527
212.6178 |  | 213.2528
213.4573 |  | 213.2528
215.8731 |  | 213.2533
211.9992 |  | 213.2533
214.5958 |  | 213.2537
214.7725 |  | 213.254
212.7946 |  | 213.2546
213.7177 |  | 213.2559
212.3748 |  | 213.2563
213.2533 |  | 213.2568
213.3686 |  | 213.257
212.4903 |  | 213.2572
214.0954 |  | 213.2573
213.8904 |  | 213.2575
212.0935 |  | 213.2578
213.2413 |  | 213.2586
214.3437 |  | 213.2602
213.1537 |  | 213.2603
214.1265 |  | 213.2606
212.846 |  | 213.2609
212.8026 |  | 213.2622
212.0876 |  | 213.2626
212.8689 |  | 213.2635
214.0366 |  | 213.2635
213.58 |  | 213.2654
215.6804 |  | 213.2665
214.49 |  | 213.2666
212.3141 |  | 213.2667
211.4132 |  | 213.2669
211.5375 |  | 213.2686
213.6977 |  | 213.2695
213.6939 |  | 213.2695
212.7115 |  | 213.2699
213.75 |  | 213.2709
213.643 |  | 213.2709
213.7672 |  | 213.2728
212.9805 |  | 213.2738
214.7631 |  | 213.275
213.2654 |  | 213.2777
212.8861 |  | 213.279
214.2122 |  | 213.2792
212.1696 |  | 213.2795
215.0622 |  | 213.2803
213.4014 |  | 213.2806
212.4445 |  | 213.2807
213.1446 |  | 213.2807
212.1136 |  | 213.2814
213.2395 |  | 213.2815
213.0963 |  | 213.2819
213.4038 |  | 213.2827
211.7141 |  | 213.2831
211.9242 |  | 213.2832
213.9054 |  | 213.284
213.5174 |  | 213.2842
213.3405 |  | 213.2849
212.5284 |  | 213.2852
212.1429 |  | 213.286
214.3454 |  | 213.2863
214.3749 |  | 213.2871
214.5871 |  | 213.2883
214.8199 |  | 213.2886
212.0348 |  | 213.2886
213.5173 |  | 213.2889
212.7596 |  | 213.2891
214.0319 |  | 213.2892
212.3752 |  | 213.2899
211.0242 |  | 213.2904
211.6653 |  | 213.2907
212.5277 |  | 213.2914
212.0378 |  | 213.2917
213.9634 |  | 213.2924
213.7718 |  | 213.2934
211.8998 |  | 213.2936
213.9655 |  | 213.2941
212.8736 |  | 213.2942
212.4858 |  | 213.2943
213.6338 |  | 213.2944
213.1825 |  | 213.2949
212.4759 |  | 213.295
212.0932 |  | 213.295
213.4987 |  | 213.2951
212.545 |  | 213.2953
214.6908 |  | 213.2959
215.2442 |  | 213.2965
215.9239 |  | 213.2965
214.158 |  | 213.2966
213.2904 |  | 213.2974
212.5351 |  | 213.2976
215.3173 |  | 213.2979
212.8971 |  | 213.2982
214.5827 |  | 213.2991
213.7168 |  | 213.2998
213.4846 |  | 213.3009
214.6282 |  | 213.301
211.3563 |  | 213.3011
211.5496 |  | 213.3034
213.2728 |  | 213.3059
210.8689 |  | 213.3059
212.4509 |  | 213.306
213.7726 |  | 213.3067
212.5179 |  | 213.3068
212.3575 |  | 213.3073
212.9602 |  | 213.3083
216.0657 |  | 213.3086
213.3503 |  | 213.3091
213.3385 |  | 213.3094
213.7284 |  | 213.3099
213.4109 |  | 213.3101
214.2813 |  | 213.3105
211.0812 |  | 213.3107
211.9034 |  | 213.3109
213.7702 |  | 213.3121
211.7995 |  | 213.3123
211.8843 |  | 213.3126
213.5173 |  | 213.3134
214.6043 |  | 213.3144
214.1794 |  | 213.3145
214.5424 |  | 213.3148
213.0719 |  | 213.3148
214.3069 |  | 213.3155
212.0361 |  | 213.3157
215.161 |  | 213.3169
212.6858 |  | 213.3179
212.065 |  | 213.318
212.1777 |  | 213.3191
213.5621 |  | 213.3196
213.9092 |  | 213.3197
213.0984 |  | 213.3204
215.1556 |  | 213.3209
211.7745 |  | 213.3221
213.2095 |  | 213.3239
212.9088 |  | 213.3241
215.6667 |  | 213.3256
213.1759 |  | 213.3259
212.4219 |  | 213.3265
215.3924 |  | 213.3266
214.4662 |  | 213.3275
213.4592 |  | 213.3289
215.2059 |  | 213.3293
213.0967 |  | 213.3298
212.4265 |  | 213.3312
211.7281 |  | 213.3312
213.4835 |  | 213.3315
212.7229 |  | 213.3322
212.3311 |  | 213.3329
214.0912 |  | 213.3341
213.9607 |  | 213.335
212.3619 |  | 213.3351
213.4277 |  | 213.3351
216.3867 |  | 213.3353
213.1116 |  | 213.3356
213.2366 |  | 213.3358
212.5374 |  | 213.3358
211.3102 |  | 213.3361
215.4484 |  | 213.3363
213.5483 |  | 213.3367
213.032 |  | 213.3369
212.8698 |  | 213.3372
213.1121 |  | 213.3373
213.6613 |  | 213.3381
212.3187 |  | 213.3385
212.2649 |  | 213.339
212.2355 |  | 213.3391
212.2764 |  | 213.3391
214.5662 |  | 213.3395
214.9971 |  | 213.3403
213.3148 |  | 213.3405
212.7593 |  | 213.342
212.113 |  | 213.3423
213.0181 |  | 213.3433
214.0689 |  | 213.3435
210.8017 |  | 213.3443
213.6131 |  | 213.3445
212.6758 |  | 213.3467
213.5801 |  | 213.3475
212.8771 |  | 213.3476
213.4655 |  | 213.3477
214.5616 |  | 213.3482
212.9047 |  | 213.3485
213.1784 |  | 213.3486
212.9528 |  | 213.3487
212.7322 |  | 213.3487
211.5931 |  | 213.3503
213.7898 |  | 213.3507
214.1202 |  | 213.3515
210.8235 |  | 213.353
213.4719 |  | 213.3537
214.584 |  | 213.3546
211.3747 |  | 213.3559
212.5924 |  | 213.3567
213.996 |  | 213.3571
214.2384 |  | 213.3576
213.0213 |  | 213.3585
212.6499 |  | 213.3586
214.1217 |  | 213.3592
213.4392 |  | 213.3609
213.518 |  | 213.3613
213.5677 |  | 213.3613
214.9483 |  | 213.3614
215.0058 |  | 213.3617
211.6942 |  | 213.3618
211.9024 |  | 213.3623
213.0494 |  | 213.3624
214.1157 |  | 213.3644
213.0708 |  | 213.3662
211.8249 |  | 213.3668
215.2235 |  | 213.3676
212.1095 |  | 213.3677
214.689 |  | 213.3678
213.2622 |  | 213.368
215.6993 |  | 213.3686
214.1335 |  | 213.3687
213.9281 |  | 213.37
211.8182 |  | 213.3702
214.147 |  | 213.3708
215.4561 |  | 213.3711
214.1885 |  | 213.3714
213.8991 |  | 213.3717
212.7755 |  | 213.3718
215.195 |  | 213.3718
213.0172 |  | 213.3723
212.6145 |  | 213.3725
210.6108 |  | 213.3737
213.7873 |  | 213.3742
214.4623 |  | 213.3748
211.6501 |  | 213.375
212.8534 |  | 213.3753
211.1447 |  | 213.3758
215.1278 |  | 213.3765
212.6956 |  | 213.377
213.9484 |  | 213.3774
212.7464 |  | 213.3777
212.8582 |  | 213.3777
212.1794 |  | 213.3783
212.7027 |  | 213.3786
213.4631 |  | 213.3789
213.5855 |  | 213.3809
213.0133 |  | 213.381
213.5145 |  | 213.3811
214.5291 |  | 213.3819
213.6881 |  | 213.384
212.4041 |  | 213.3841
213.6101 |  | 213.3875
212.6436 |  | 213.3878
212.358 |  | 213.3879
213.456 |  | 213.3882
212.1707 |  | 213.3883
211.9951 |  | 213.3885
212.4913 |  | 213.3887
212.7855 |  | 213.3888
211.9782 |  | 213.3889
212.5348 |  | 213.3891
213.5266 |  | 213.3894
213.9655 |  | 213.39
214.6823 |  | 213.3904
213.9514 |  | 213.3905
213.4391 |  | 213.3907
215.477 |  | 213.3915
213.116 |  | 213.3928
213.9282 |  | 213.3934
213.2054 |  | 213.3935
211.2615 |  | 213.3937
213.964 |  | 213.3953
212.4624 |  | 213.3967
213.4805 |  | 213.3968
211.6743 |  | 213.3981
214.4471 |  | 213.3983
211.6836 |  | 213.3989
214.3103 |  | 213.399
214.6606 |  | 213.4014
213.3882 |  | 213.4019
213.4132 |  | 213.4021
212.8104 |  | 213.4022
215.656 |  | 213.4025
212.6502 |  | 213.4027
212.1982 |  | 213.403
214.463 |  | 213.4033
212.2125 |  | 213.4036
212.6473 |  | 213.4038
214.168 |  | 213.4059
212.9686 |  | 213.406
214.0913 |  | 213.4064
214.4886 |  | 213.4064
214.0208 |  | 213.4066
213.95 |  | 213.4071
214.7755 |  | 213.408
213.0891 |  | 213.4085
212.9944 |  | 213.4087
212.9265 |  | 213.4088
212.6839 |  | 213.4098
214.6025 |  | 213.4109
211.9843 |  | 213.4112
212.1101 |  | 213.4113
214.2421 |  | 213.4114
213.7173 |  | 213.4122
213.5656 |  | 213.4126
212.7633 |  | 213.4132
215.0206 |  | 213.4144
214.5972 |  | 213.4152
212.4249 |  | 213.4157
211.8103 |  | 213.4183
213.1768 |  | 213.4192
213.9564 |  | 213.42
214.7162 |  | 213.422
210.9489 |  | 213.4224
214.4544 |  | 213.4226
213.8049 |  | 213.4234
212.8461 |  | 213.424
213.158 |  | 213.424
213.7772 |  | 213.4248
214.1786 |  | 213.4249
212.6536 |  | 213.4277
214.271 |  | 213.4279
212.7943 |  | 213.4286
214.5119 |  | 213.4299
212.6062 |  | 213.43
211.7124 |  | 213.4306
212.8005 |  | 213.4307
213.1322 |  | 213.4307
211.7367 |  | 213.4321
215.884 |  | 213.4325
213.7371 |  | 213.4328
213.8046 |  | 213.4334
213.6219 |  | 213.4353
213.2472 |  | 213.4358
212.1076 |  | 213.4391
211.8917 |  | 213.4391
212.1724 |  | 213.4392
212.8901 |  | 213.4394
212.4694 |  | 213.4395
214.5083 |  | 213.4396
214.0676 |  | 213.4398
214.7669 |  | 213.4398
214.0969 |  | 213.4398
213.3678 |  | 213.4401
213.9817 |  | 213.4402
214.5125 |  | 213.4411
213.817 |  | 213.4418
213.4751 |  | 213.4421
214.8887 |  | 213.4422
211.2037 |  | 213.4424
211.9632 |  | 213.4424
214.8407 |  | 213.4425
213.1621 |  | 213.4428
212.8361 |  | 213.4429
212.5921 |  | 213.4431
213.8017 |  | 213.4439
214.4724 |  | 213.4442
213.8028 |  | 213.4449
214.6729 |  | 213.4452
212.3747 |  | 213.4452
211.2679 |  | 213.4453
214.4522 |  | 213.4458
213.3576 |  | 213.4467
212.777 |  | 213.4475
213.7917 |  | 213.4482
213.5749 |  | 213.4484
215.1119 |  | 213.4487
213.8518 |  | 213.4489
212.0425 |  | 213.4499
212.1697 |  | 213.4507
215.2377 |  | 213.4519
212.8366 |  | 213.4519
214.641 |  | 213.4525
213.3742 |  | 213.4527
212.5289 |  | 213.4535
212.5544 |  | 213.4539
213.5258 |  | 213.4547
213.295 |  | 213.4552
212.6964 |  | 213.4554
211.9849 |  | 213.4557
213.2953 |  | 213.4558
212.7774 |  | 213.456
214.5925 |  | 213.4564
213.0699 |  | 213.4565
211.692 |  | 213.4573
215.3221 |  | 213.4573
212.9565 |  | 213.4579
212.3206 |  | 213.4582
212.6838 |  | 213.4589
212.2442 |  | 213.4592
212.7652 |  | 213.4594
213.0713 |  | 213.4595
212.8273 |  | 213.4597
213.0256 |  | 213.4613
213.805 |  | 213.4622
213.2173 |  | 213.4623
214.3636 |  | 213.4628
212.5007 |  | 213.4631
211.9157 |  | 213.4631
214.7222 |  | 213.4638
212.9997 |  | 213.4644
212.6902 |  | 213.4655
211.9279 |  | 213.4657
213.3809 |  | 213.4687
212.782 |  | 213.4695
213.632 |  | 213.4698
212.729 |  | 213.4704
212.0577 |  | 213.4705
213.3894 |  | 213.4707
214.6911 |  | 213.4709
214.4602 |  | 213.471
215.3774 |  | 213.4716
214.646 |  | 213.4718
213.101 |  | 213.4719
213.9751 |  | 213.4721
214.6765 |  | 213.4728
212.5013 |  | 213.4728
213.3623 |  | 213.4732
213.5691 |  | 213.4749
213.877 |  | 213.4749
215.2706 |  | 213.4751
211.2785 |  | 213.4758
211.6825 |  | 213.476
212.5172 |  | 213.476
214.8132 |  | 213.476
214.9712 |  | 213.477
212.3272 |  | 213.4771
211.7217 |  | 213.4773
212.3372 |  | 213.4781
212.2423 |  | 213.4784
212.7216 |  | 213.4787
213.4064 |  | 213.4796
213.7953 |  | 213.4805
212.4123 |  | 213.4811
213.0459 |  | 213.4819
212.7581 |  | 213.4828
214.3389 |  | 213.4835
213.2111 |  | 213.4846
212.2497 |  | 213.4846
213.4402 |  | 213.4848
212.9556 |  | 213.4849
213.9166 |  | 213.4852
214.8147 |  | 213.4868
213.6997 |  | 213.488
211.7902 |  | 213.4883
212.0414 |  | 213.49
214.3754 |  | 213.4906
214.5088 |  | 213.4909
212.5074 |  | 213.4922
212.614 |  | 213.4923
213.4424 |  | 213.4937
212.1783 |  | 213.4937
212.3515 |  | 213.4942
214.2706 |  | 213.4947
213.2407 |  | 213.4948
210.8673 |  | 213.495
213.6323 |  | 213.4971
212.1161 |  | 213.4983
212.712 |  | 213.4987
213.8998 |  | 213.4996
212.3604 |  | 213.4997
215.3453 |  | 213.4999
211.2347 |  | 213.4999
213.8146 |  | 213.5004
213.2519 |  | 213.5009
213.7854 |  | 213.5011
213.2819 |  | 213.5025
212.6918 |  | 213.5029
215.6741 |  | 213.503
212.3532 |  | 213.5031
214.6398 |  | 213.5039
212.9992 |  | 213.5039
212.0011 |  | 213.5043
212.8484 |  | 213.5045
211.8006 |  | 213.5045
214.2365 |  | 213.5049
212.0287 |  | 213.5058
213.5732 |  | 213.507
213.0297 |  | 213.5076
212.6296 |  | 213.5082
213.0571 |  | 213.5095
213.2269 |  | 213.5117
212.7557 |  | 213.5132
214.7224 |  | 213.5136
212.7975 |  | 213.5138
212.1387 |  | 213.5143
212.7769 |  | 213.5145
211.123 |  | 213.5147
213.9091 |  | 213.5148
212.8659 |  | 213.5148
212.3483 |  | 213.515
213.4353 |  | 213.515
212.6616 |  | 213.5151
213.1494 |  | 213.5153
212.132 |  | 213.516
213.8996 |  | 213.516
213.1879 |  | 213.5165
213.4564 |  | 213.5168
212.349 |  | 213.5173
212.0457 |  | 213.5173
215.1196 |  | 213.5174
211.7888 |  | 213.518
213.7818 |  | 213.5182
212.7867 |  | 213.5193
211.6029 |  | 213.5195
212.1637 |  | 213.5221
213.4758 |  | 213.5222
211.3934 |  | 213.5225
212.6599 |  | 213.5227
211.722 |  | 213.523
213.9748 |  | 213.5231
212.8653 |  | 213.5235
213.8205 |  | 213.5258
212.0146 |  | 213.5258
213.1435 |  | 213.5266
213.1322 |  | 213.5266
213.6064 |  | 213.5271
211.184 |  | 213.5271
214.4046 |  | 213.5277
212.3675 |  | 213.5278
214.1619 |  | 213.5278
215.4144 |  | 213.5318
212.0566 |  | 213.5328
212.6962 |  | 213.5329
214.4117 |  | 213.5332
213.4299 |  | 213.5349
215.0324 |  | 213.5353
213.8015 |  | 213.5353
215.02 |  | 213.5357
212.5354 |  | 213.536
214.4627 |  | 213.5366
212.0883 |  | 213.5373
214.1049 |  | 213.5378
212.7951 |  | 213.5381
213.9947 |  | 213.5386
211.4737 |  | 213.5387
214.3096 |  | 213.5399
212.8101 |  | 213.5399
211.5209 |  | 213.5405
212.4474 |  | 213.5414
211.6699 |  | 213.5418
214.8954 |  | 213.5423
210.9366 |  | 213.5428
213.0427 |  | 213.544
213.1631 |  | 213.5448
214.2292 |  | 213.5465
211.8649 |  | 213.5466
212.2075 |  | 213.547
213.7668 |  | 213.5471
213.4657 |  | 213.5479
213.1722 |  | 213.5483
212.6344 |  | 213.5492
213.6722 |  | 213.5492
214.3828 |  | 213.5507
213.5258 |  | 213.5515
212.9524 |  | 213.5524
213.2871 |  | 213.5525
213.3928 |  | 213.553
214.3465 |  | 213.5531
213.1056 |  | 213.5532
213.0694 |  | 213.5536
215.2453 |  | 213.5545
211.8148 |  | 213.556
214.844 |  | 213.5569
212.9883 |  | 213.5569
213.6533 |  | 213.557
212.9703 |  | 213.5571
212.6676 |  | 213.5571
213.6824 |  | 213.5574
214.9232 |  | 213.5576
213.0869 |  | 213.5577
214.3051 |  | 213.558
213.9552 |  | 213.5589
212.7846 |  | 213.559
211.2516 |  | 213.559
212.9748 |  | 213.5598
212.2836 |  | 213.56
212.3682 |  | 213.5616
213.5332 |  | 213.5621
213.0831 |  | 213.5623
211.8533 |  | 213.5625
211.8313 |  | 213.5641
212.6646 |  | 213.5643
214.5211 |  | 213.5647
214.0051 |  | 213.5656
212.6646 |  | 213.5659
213.771 |  | 213.566
211.426 |  | 213.5665
213.616 |  | 213.5666
213.0035 |  | 213.5672
213.4224 |  | 213.5675
211.9743 |  | 213.5677
214.8408 |  | 213.5677
214.9228 |  | 213.5682
212.9217 |  | 213.5691
214.8782 |  | 213.5691
213.3875 |  | 213.5692
210.8998 |  | 213.5694
212.5907 |  | 213.5697
211.7885 |  | 213.5698
212.377 |  | 213.5716
213.3059 |  | 213.5732
213.1815 |  | 213.5749
214.0659 |  | 213.5753
214.8122 |  | 213.5754
212.9778 |  | 213.5763
213.724 |  | 213.5777
215.498 |  | 213.5778
214.3204 |  | 213.5783
212.8495 |  | 213.5796
214.5892 |  | 213.5796
212.0008 |  | 213.58
213.9396 |  | 213.5801
214.162 |  | 213.5804
213.7323 |  | 213.5806
214.1782 |  | 213.5816
211.8301 |  | 213.5829
212.8238 |  | 213.5836
212.6698 |  | 213.5841
213.4422 |  | 213.5855
213.3811 |  | 213.5856
213.8504 |  | 213.5889
213.1456 |  | 213.5891
212.7938 |  | 213.5906
214.2437 |  | 213.5911
211.0018 |  | 213.5915
214.3358 |  | 213.5918
213.2666 |  | 213.592
212.2165 |  | 213.5924
212.9998 |  | 213.5928
215.2385 |  | 213.5955
211.1284 |  | 213.5956
212.0734 |  | 213.5956
213.4937 |  | 213.5966
213.477 |  | 213.597
213.8717 |  | 213.5987
213.7525 |  | 213.6013
215.1181 |  | 213.6017
213.424 |  | 213.6031
211.5075 |  | 213.6037
213.8796 |  | 213.6055
212.0556 |  | 213.6057
212.6338 |  | 213.6064
212.8057 |  | 213.6068
213.5675 |  | 213.6073
213.9277 |  | 213.6076
214.0895 |  | 213.6087
213.6115 |  | 213.6093
213.7762 |  | 213.6101
215.1731 |  | 213.6115
212.304 |  | 213.6115
213.2225 |  | 213.6131
210.6874 |  | 213.6132
214.0339 |  | 213.6147
213.8161 |  | 213.6147
213.5697 |  | 213.6149
214.5553 |  | 213.616
214.7504 |  | 213.6173
214.1439 |  | 213.6177
213.5148 |  | 213.6189
213.7408 |  | 213.6191
214.0873 |  | 213.6193
214.4093 |  | 213.6197
213.2803 |  | 213.6208
215.2571 |  | 213.6219
212.7114 |  | 213.6226
215.0081 |  | 213.6228
212.4462 |  | 213.6232
213.6447 |  | 213.6239
213.9575 |  | 213.6239
214.2427 |  | 213.625
214.6447 |  | 213.6254
214.637 |  | 213.6274
213.1197 |  | 213.6274
212.4564 |  | 213.6278
213.723 |  | 213.628
212.8239 |  | 213.6298
214.5567 |  | 213.6303
213.5195 |  | 213.6308
213.9483 |  | 213.6318
213.0404 |  | 213.632
212.706 |  | 213.6323
210.8197 |  | 213.6329
211.3649 |  | 213.6329
212.2782 |  | 213.6329
213.7942 |  | 213.633
212.4701 |  | 213.6333
213.1646 |  | 213.6338
212.57 |  | 213.634
215.063 |  | 213.6342
213.8573 |  | 213.6349
213.6303 |  | 213.6351
213.2024 |  | 213.6373
211.2733 |  | 213.6381
212.5519 |  | 213.6383
213.0043 |  | 213.6392
214.5946 |  | 213.6422
215.6717 |  | 213.6425
212.4663 |  | 213.643
215.4049 |  | 213.6431
214.2269 |  | 213.6434
214.3703 |  | 213.6438
213.3718 |  | 213.6441
214.4275 |  | 213.6442
213.808 |  | 213.6447
213.4122 |  | 213.6461
214.428 |  | 213.6463
214.9182 |  | 213.6468
215.0225 |  | 213.6473
211.7159 |  | 213.6475
212.7316 |  | 213.6484
211.6845 |  | 213.6487
212.8003 |  | 213.6503
213.3786 |  | 213.6509
214.0287 |  | 213.651
214.2875 |  | 213.6521
212.8584 |  | 213.6522
213.9871 |  | 213.6533
214.1397 |  | 213.6533
213.0008 |  | 213.6537
212.138 |  | 213.6545
213.4126 |  | 213.6546
212.3805 |  | 213.6551
213.7449 |  | 213.6555
214.505 |  | 213.6562
214.1563 |  | 213.6572
213.275 |  | 213.6586
215.3359 |  | 213.659
212.4286 |  | 213.661
213.5694 |  | 213.6613
213.7119 |  | 213.6622
212.4158 |  | 213.6631
214.4752 |  | 213.6633
211.6633 |  | 213.6639
213.2792 |  | 213.6639
213.4183 |  | 213.665
212.9071 |  | 213.6652
212.1824 |  | 213.6653
212.8199 |  | 213.6656
214.2671 |  | 213.67
211.83 |  | 213.6707
213.1506 |  | 213.6708
213.3717 |  | 213.6712
212.0696 |  | 213.672
215.2804 |  | 213.6722
212.7101 |  | 213.6723
212.4325 |  | 213.6725
212.2996 |  | 213.6727
211.6709 |  | 213.6732
215.5608 |  | 213.6733
213.6732 |  | 213.6739
213.4418 |  | 213.6741
215.5723 |  | 213.6742
213.1596 |  | 213.675
216.1854 |  | 213.6753
212.5593 |  | 213.6774
212.649 |  | 213.6776
215.0944 |  | 213.6792
214.3076 |  | 213.6807
211.9844 |  | 213.6824
213.3644 |  | 213.6834
212.2508 |  | 213.6835
212.9979 |  | 213.684
214.2401 |  | 213.6861
213.2944 |  | 213.6873
212.8342 |  | 213.688
211.983 |  | 213.6881
212.3217 |  | 213.6881
212.8118 |  | 213.6897
214.5874 |  | 213.6939
211.7796 |  | 213.6939
211.9582 |  | 213.6944
213.2899 |  | 213.6956
212.2415 |  | 213.696
214.2519 |  | 213.6963
211.827 |  | 213.697
215.6409 |  | 213.6977
213.2513 |  | 213.6984
212.1132 |  | 213.6997
212.0133 |  | 213.7002
212.9852 |  | 213.701
213.5039 |  | 213.7011
212.9711 |  | 213.7013
214.9576 |  | 213.7014
211.7309 |  | 213.7025
212.9169 |  | 213.7028
214.522 |  | 213.7037
212.1591 |  | 213.7042
213.133 |  | 213.7064
215.9572 |  | 213.7064
214.3713 |  | 213.7066
214.5122 |  | 213.7071
213.3353 |  | 213.7075
213.2346 |  | 213.7076
213.6551 |  | 213.7084
213.7357 |  | 213.7092
212.2355 |  | 213.7094
214.6687 |  | 213.7097
211.7477 |  | 213.71
212.9497 |  | 213.7108
212.9136 |  | 213.7119
212.8912 |  | 213.713
214.5426 |  | 213.7133
213.7574 |  | 213.7159
214.8571 |  | 213.7168
213.3487 |  | 213.7173
211.3084 |  | 213.7177
213.5387 |  | 213.7177
214.9997 |  | 213.719
214.6512 |  | 213.7191
214.2571 |  | 213.7197
213.4558 |  | 213.72
211.8052 |  | 213.7201
212.0555 |  | 213.7203
213.2229 |  | 213.7205
212.5329 |  | 213.7216
214.5673 |  | 213.7217
213.544 |  | 213.723
215.5302 |  | 213.724
213.251 |  | 213.7252
214.2698 |  | 213.7259
213.6572 |  | 213.7284
214.5657 |  | 213.7311
210.6423 |  | 213.7321
214.5478 |  | 213.7323
213.207 |  | 213.7324
213.5778 |  | 213.7324
213.4539 |  | 213.7329
214.6033 |  | 213.7331
214.9424 |  | 213.7331
213.507 |  | 213.7341
213.9765 |  | 213.7347
212.5967 |  | 213.735
213.852 |  | 213.7357
214.3989 |  | 213.7359
212.9773 |  | 213.7369
214.4932 |  | 213.7371
214.9386 |  | 213.7382
212.9159 |  | 213.7384
211.852 |  | 213.7389
213.6509 |  | 213.7394
212.3844 |  | 213.7404
213.4828 |  | 213.7407
215.8389 |  | 213.7408
213.476 |  | 213.7417
213.7989 |  | 213.7418
214.4598 |  | 213.7425
214.2356 |  | 213.7425
213.592 |  | 213.7428
213.3934 |  | 213.7437
212.2161 |  | 213.7443
213.5466 |  | 213.7444
210.9352 |  | 213.7449
213.6191 |  | 213.7451
212.1694 |  | 213.7455
213.4452 |  | 213.746
214.6951 |  | 213.7462
211.4105 |  | 213.7468
214.2636 |  | 213.7482
214.196 |  | 213.7487
213.7686 |  | 213.7489
211.5401 |  | 213.7495
214.3516 |  | 213.75
213.2709 |  | 213.7506
213.3618 |  | 213.7513
213.2525 |  | 213.7515
213.5532 |  | 213.7518
215.0472 |  | 213.7521
213.254 |  | 213.7525
212.508 |  | 213.753
214.0783 |  | 213.7546
211.4375 |  | 213.7548
214.6843 |  | 213.7551
215.079 |  | 213.7554
213.015 |  | 213.7567
213.8817 |  | 213.7571
213.8308 |  | 213.7574
214.8625 |  | 213.7581
213.7025 |  | 213.7588
214.0151 |  | 213.7594
213.0914 |  | 213.7602
212.2915 |  | 213.7611
215.3462 |  | 213.7628
214.0147 |  | 213.7628
215.0991 |  | 213.7629
213.4157 |  | 213.7638
214.7521 |  | 213.7638
213.3123 |  | 213.7639
212.4007 |  | 213.7642
213.9531 |  | 213.7662
213.1779 |  | 213.7668
213.158 |  | 213.7672
214.1097 |  | 213.7681
214.3668 |  | 213.7683
213.6017 |  | 213.7686
213.286 |  | 213.7702
214.7573 |  | 213.7705
213.5147 |  | 213.771
213.5278 |  | 213.7718
213.5138 |  | 213.7721
211.604 |  | 213.7724
213.3887 |  | 213.7726
213.8976 |  | 213.7729
213.7546 |  | 213.7733
212.5865 |  | 213.7739
215.3271 |  | 213.7744
213.5182 |  | 213.7747
213.3937 |  | 213.7761
214.2193 |  | 213.7762
212.8935 |  | 213.7763
211.9948 |  | 213.7765
213.9479 |  | 213.7772
215.5912 |  | 213.7778
214.834 |  | 213.7818
212.7863 |  | 213.7827
212.4138 |  | 213.7839
213.5049 |  | 213.7844
212.3801 |  | 213.7854
211.6696 |  | 213.7855
214.3498 |  | 213.786
214.0987 |  | 213.786
213.7425 |  | 213.7869
213.4192 |  | 213.7873
212.3255 |  | 213.7873
213.6723 |  | 213.7876
212.9877 |  | 213.7878
210.6288 |  | 213.7879
213.4557 |  | 213.7896
213.8623 |  | 213.7898
214.2346 |  | 213.7917
214.3666 |  | 213.7917
212.9986 |  | 213.793
213.2337 |  | 213.7931
210.9089 |  | 213.7931
213.5698 |  | 213.7942
214.1066 |  | 213.7943
214.3782 |  | 213.7949
212.3056 |  | 213.7953
212.1345 |  | 213.7983
211.8725 |  | 213.7989
213.5153 |  | 213.7994
214.857 |  | 213.8005
214.9934 |  | 213.8011
213.1021 |  | 213.8015
213.1496 |  | 213.8017
210.5647 |  | 213.8017
212.7333 |  | 213.8018
213.2282 |  | 213.8028
212.1923 |  | 213.8029
212.5579 |  | 213.8032
211.9283 |  | 213.804
213.5386 |  | 213.8046
214.2787 |  | 213.8049
216.3374 |  | 213.805
211.9122 |  | 213.808
214.1686 |  | 213.8085
211.0559 |  | 213.8108
212.1644 |  | 213.8109
213.4909 |  | 213.8112
215.1085 |  | 213.8118
213.8779 |  | 213.8121
213.5589 |  | 213.8145
215.7371 |  | 213.8146
211.6264 |  | 213.8148
211.7231 |  | 213.8157
212.1954 |  | 213.8161
215.563 |  | 213.8161
214.6349 |  | 213.817
213.3204 |  | 213.8183
211.2873 |  | 213.8205
213.0045 |  | 213.8213
213.7567 |  | 213.8216
213.6037 |  | 213.8216
213.3841 |  | 213.8225
214.2594 |  | 213.8229
213.0991 |  | 213.8258
214.5972 |  | 213.826
213.161 |  | 213.8267
213.0056 |  | 213.8267
211.4978 |  | 213.8269
212.5715 |  | 213.8274
214.7014 |  | 213.8283
214.6196 |  | 213.829
212.9743 |  | 213.829
214.2342 |  | 213.8296
215.644 |  | 213.8308
213.6254 |  | 213.8318
213.5418 |  | 213.8321
214.5627 |  | 213.8328
213.8157 |  | 213.8341
212.3759 |  | 213.8346
214.4553 |  | 213.8361
212.5958 |  | 213.8367
213.0145 |  | 213.8373
213.4439 |  | 213.8375
212.6591 |  | 213.8385
215.0368 |  | 213.8386
214.4693 |  | 213.8387
212.4286 |  | 213.8393
213.4565 |  | 213.84
212.296 |  | 213.8418
213.3145 |  | 213.8419
213.7554 |  | 213.8433
214.0899 |  | 213.8437
214.5751 |  | 213.8463
211.3774 |  | 213.8477
213.4597 |  | 213.8478
211.6689 |  | 213.8504
212.508 |  | 213.8508
214.0375 |  | 213.8517
213.7331 |  | 213.8517
212.3206 |  | 213.8518
214.4612 |  | 213.852
213.4507 |  | 213.8525
212.6974 |  | 213.8528
214.0494 |  | 213.8534
213.8229 |  | 213.854
210.0217 |  | 213.8541
214.4497 |  | 213.8544
212.7504 |  | 213.8553
212.2102 |  | 213.8557
214.7452 |  | 213.8558
213.4027 |  | 213.8563
214.34 |  | 213.8567
212.253 |  | 213.8573
213.222 |  | 213.8581
212.777 |  | 213.8606
212.6875 |  | 213.8613
213.9366 |  | 213.8623
212.7727 |  | 213.8633
215.6792 |  | 213.864
213.2003 |  | 213.8653
211.6022 |  | 213.8654
213.6807 |  | 213.8665
212.9312 |  | 213.8676
212.631 |  | 213.8681
214.2311 |  | 213.8707
212.3521 |  | 213.871
215.0248 |  | 213.8716
212.9556 |  | 213.8717
213.2777 |  | 213.873
213.3329 |  | 213.8732
213.4547 |  | 213.8739
214.6247 |  | 213.8739
213.3381 |  | 213.8742
214.2135 |  | 213.8746
212.0944 |  | 213.8749
213.2388 |  | 213.8762
214.8888 |  | 213.8768
211.7605 |  | 213.877
215.4602 |  | 213.877
214.4855 |  | 213.8771
212.8512 |  | 213.8778
213.8328 |  | 213.8779
213.4449 |  | 213.8784
211.7323 |  | 213.8786
211.5101 |  | 213.8788
212.3771 |  | 213.8796
214.0536 |  | 213.8798
214.8887 |  | 213.8804
212.8018 |  | 213.8805
214.822 |  | 213.8808
212.2046 |  | 213.8817
213.0014 |  | 213.8824
214.0684 |  | 213.8828
214.7778 |  | 213.8834
212.4074 |  | 213.8834
212.8682 |  | 213.8841
211.4524 |  | 213.8847
213.2018 |  | 213.8849
211.6688 |  | 213.8853
210.6997 |  | 213.8854
213.1896 |  | 213.8858
213.8933 |  | 213.8858
212.3021 |  | 213.8863
213.5043 |  | 213.8868
215.2329 |  | 213.887
212.451 |  | 213.8904
212.2137 |  | 213.8912
212.3038 |  | 213.8916
212.7852 |  | 213.8925
212.0871 |  | 213.8927
213.8321 |  | 213.8931
215.479 |  | 213.8933
212.3837 |  | 213.8943
211.6241 |  | 213.8962
213.9587 |  | 213.8966
213.7983 |  | 213.8967
213.3774 |  | 213.8974
213.8974 |  | 213.8974
214.4079 |  | 213.8976
213.295 |  | 213.8982
212.9474 |  | 213.8982
214.9171 |  | 213.8984
213.8676 |  | 213.8986
212.7565 |  | 213.899
213.871 |  | 213.8991
214.3406 |  | 213.8995
212.3426 |  | 213.8996
213.4771 |  | 213.8998
214.5046 |  | 213.9001
213.3891 |  | 213.9016
213.9589 |  | 213.9023
210.7055 |  | 213.9031
211.6687 |  | 213.9031
214.2061 |  | 213.9031
210.5495 |  | 213.9041
213.7013 |  | 213.9041
210.5978 |  | 213.9046
212.926 |  | 213.9054
213.3091 |  | 213.9057
214.4972 |  | 213.9066
213.2635 |  | 213.9077
213.019 |  | 213.9078
214.7227 |  | 213.9091
213.5148 |  | 213.9092
211.2148 |  | 213.9107
212.1157 |  | 213.9116
213.0366 |  | 213.9124
213.2133 |  | 213.9127
212.9655 |  | 213.913
213.4705 |  | 213.9137
213.4589 |  | 213.9154
212.485 |  | 213.9157
213.3289 |  | 213.916
211.8069 |  | 213.9164
212.9786 |  | 213.9166
214.8553 |  | 213.9202
211.9694 |  | 213.9202
211.5588 |  | 213.9202
213.6173 |  | 213.9214
212.416 |  | 213.923
214.7398 |  | 213.9238
211.2093 |  | 213.9244
212.3597 |  | 213.9244
213.1194 |  | 213.9257
211.9243 |  | 213.9275
211.5142 |  | 213.9275
213.5576 |  | 213.9277
213.1843 |  | 213.9281
212.9703 |  | 213.9282
212.049 |  | 213.9294
214.2093 |  | 213.9332
212.2077 |  | 213.9346
211.5786 |  | 213.9349
211.8858 |  | 213.9355
214.5943 |  | 213.9358
212.539 |  | 213.9361
213.873 |  | 213.9366
212.4735 |  | 213.9374
213.2507 |  | 213.9377
212.8109 |  | 213.9395
213.6963 |  | 213.9396
213.0687 |  | 213.9396
212.5914 |  | 213.9407
212.1065 |  | 213.9408
213.1416 |  | 213.9413
213.8804 |  | 213.9413
213.0818 |  | 213.9427
213.1282 |  | 213.9429
212.397 |  | 213.9439
215.3508 |  | 213.944
213.6438 |  | 213.9451
213.9757 |  | 213.9461
214.3757 |  | 213.9474
214.7364 |  | 213.9477
214.7671 |  | 213.9478
211.1223 |  | 213.9479
212.7289 |  | 213.9483
212.6338 |  | 213.9484
214.0174 |  | 213.9493
213.4535 |  | 213.95
213.753 |  | 213.9502
213.9692 |  | 213.9508
214.5234 |  | 213.9514
214.3667 |  | 213.9531
213.9645 |  | 213.9533
215.4888 |  | 213.9533
212.7233 |  | 213.9534
211.3616 |  | 213.9552
213.144 |  | 213.9556
212.7741 |  | 213.9557
213.2196 |  | 213.9564
215.0845 |  | 213.9575
213.3567 |  | 213.9577
214.0488 |  | 213.958
212.3514 |  | 213.9581
214.6014 |  | 213.9587
213.6189 |  | 213.9588
214.7379 |  | 213.9589
214.3042 |  | 213.96
212.595 |  | 213.9603
216.5549 |  | 213.9607
214.2526 |  | 213.9634
212.6657 |  | 213.9634
212.4782 |  | 213.9637
214.5556 |  | 213.964
214.2973 |  | 213.9645
214.6471 |  | 213.9645
213.0744 |  | 213.9655
212.581 |  | 213.9655
213.9888 |  | 213.9656
211.9241 |  | 213.9658
214.2192 |  | 213.9659
213.284 |  | 213.9672
213.3585 |  | 213.9686
214.7513 |  | 213.9688
213.8517 |  | 213.9691
212.546 |  | 213.9692
212.577 |  | 213.9692
213.3907 |  | 213.9693
214.1694 |  | 213.9707
212.4813 |  | 213.9713
213.1338 |  | 213.9724
214.3737 |  | 213.9727
214.3997 |  | 213.9734
214.2256 |  | 213.9736
214.7067 |  | 213.9737
212.0474 |  | 213.9748
212.7877 |  | 213.975
213.3221 |  | 213.9751
213.0687 |  | 213.9757
213.4088 |  | 213.9765
213.1275 |  | 213.9788
213.2695 |  | 213.9789
212.9552 |  | 213.9798
213.651 |  | 213.9811
213.5231 |  | 213.9817
212.2587 |  | 213.9825
212.0146 |  | 213.9833
214.015 |  | 213.9834
210.3757 |  | 213.9843
215.0835 |  | 213.9853
213.6562 |  | 213.9853
212.5882 |  | 213.9856
212.7172 |  | 213.9871
213.4554 |  | 213.9888
211.6442 |  | 213.989
214.3288 |  | 213.9895
212.8271 |  | 213.9908
212.047 |  | 213.9916
214.7324 |  | 213.9939
213.553 |  | 213.9947
212.4173 |  | 213.9958
214.3614 |  | 213.996
214.7871 |  | 213.9964
213.8108 |  | 213.9968
215.7457 |  | 213.9987
212.6213 |  | 213.9992
213.3356 |  | 214.0006
211.4373 |  | 214.0038
213.5448 |  | 214.004
211.7182 |  | 214.004
214.0615 |  | 214.0047
215.1836 |  | 214.0051
212.321 |  | 214.0051
213.3883 |  | 214.0056
213.4784 |  | 214.0057
211.5846 |  | 214.0059
214.4987 |  | 214.0074
214.3178 |  | 214.0087
213.8858 |  | 214.0091
213.212 |  | 214.0095
213.5222 |  | 214.0103
213.7311 |  | 214.0109
213.4732 |  | 214.0112
213.9164 |  | 214.0128
212.997 |  | 214.0147
211.6538 |  | 214.015
211.8953 |  | 214.015
212.1095 |  | 214.0151
214.1469 |  | 214.0159
212.1617 |  | 214.0161
213.06 |  | 214.0166
214.178 |  | 214.0172
212.574 |  | 214.0174
213.3443 |  | 214.019
212.7457 |  | 214.0196
212.7276 |  | 214.0208
213.4519 |  | 214.021
212.5053 |  | 214.0217
213.5004 |  | 214.0227
212.6606 |  | 214.0232
214.9023 |  | 214.0241
212.847 |  | 214.0244
212.1394 |  | 214.0256
211.8053 |  | 214.0287
212.7668 |  | 214.0296
212.5408 |  | 214.0299
212.7504 |  | 214.0304
213.3126 |  | 214.0307
212.0413 |  | 214.0311
212.7002 |  | 214.0316
213.8216 |  | 214.0319
212.4823 |  | 214.0322
212.5711 |  | 214.0339
211.9396 |  | 214.0342
212.9245 |  | 214.0346
212.5769 |  | 214.0366
214.7455 |  | 214.0375
213.3702 |  | 214.0377
213.0518 |  | 214.0394
213.4595 |  | 214.0398
213.8633 |  | 214.0426
212.0643 |  | 214.043
213.0136 |  | 214.0436
214.2212 |  | 214.0443
211.3523 |  | 214.0446
211.3431 |  | 214.0452
211.6164 |  | 214.0452
214.059 |  | 214.0467
212.055 |  | 214.0478
213.7028 |  | 214.0486
213.7515 |  | 214.0488
215.6686 |  | 214.0494
212.7715 |  | 214.0497
214.0566 |  | 214.0516
214.3407 |  | 214.0522
213.8868 |  | 214.0534
215.5114 |  | 214.0536
212.9194 |  | 214.0539
214.3944 |  | 214.0541
213.1297 |  | 214.0544
215.4585 |  | 214.0546
216.208 |  | 214.0549
214.3569 |  | 214.0553
212.9194 |  | 214.0559
213.5836 |  | 214.0563
213.5623 |  | 214.0566
213.7462 |  | 214.059
212.4896 |  | 214.0605
214.3286 |  | 214.0614
214.3012 |  | 214.0615
214.0867 |  | 214.0618
213.6545 |  | 214.0627
213.0617 |  | 214.0629
215.0349 |  | 214.0652
213.2211 |  | 214.0659
213.7259 |  | 214.0663
213.8863 |  | 214.0665
213.1291 |  | 214.0676
211.8266 |  | 214.0683
214.021 |  | 214.0684
212.8601 |  | 214.0689
214.6501 |  | 214.0696
213.2814 |  | 214.0708
213.9856 |  | 214.0721
212.6205 |  | 214.0725
213.2082 |  | 214.0728
213.2979 |  | 214.0734
214.7789 |  | 214.0735
212.1053 |  | 214.0738
213.335 |  | 214.0762
212.4496 |  | 214.0764
213.2982 |  | 214.0781
212.2277 |  | 214.0783
213.3662 |  | 214.0818
212.6018 |  | 214.0839
212.9532 |  | 214.0849
214.7809 |  | 214.0861
211.3633 |  | 214.0867
214.9763 |  | 214.0867
214.4065 |  | 214.0873
212.6207 |  | 214.088
212.9137 |  | 214.0894
213.7506 |  | 214.0895
212.3507 |  | 214.0898
212.9183 |  | 214.0899
213.1845 |  | 214.0907
213.8385 |  | 214.0911
213.2917 |  | 214.0912
214.4549 |  | 214.0913
213.9461 |  | 214.0913
214.441 |  | 214.0913
215.2829 |  | 214.0923
210.9782 |  | 214.0925
214.1876 |  | 214.0927
214.963 |  | 214.0947
213.7177 |  | 214.0951
213.5928 |  | 214.0954
214.4872 |  | 214.0955
212.3567 |  | 214.0957
213.0362 |  | 214.0969
213.05 |  | 214.0969
212.9951 |  | 214.0987
214.6785 |  | 214.0997
213.9656 |  | 214.1004
214.7784 |  | 214.101
212.9042 |  | 214.1014
214.6117 |  | 214.1027
213.0266 |  | 214.104
215.0047 |  | 214.1047
213.6441 |  | 214.1047
213.0943 |  | 214.1049
213.4398 |  | 214.105
214.3055 |  | 214.1051
212.9514 |  | 214.1051
213.8974 |  | 214.1066
213.1699 |  | 214.107
213.2033 |  | 214.1076
213.6776 |  | 214.1078
214.8543 |  | 214.1097
215.7154 |  | 214.1097
213.5076 |  | 214.1103
212.8862 |  | 214.1122
214.1051 |  | 214.1122
214.1947 |  | 214.1125
213.5956 |  | 214.1136
213.6461 |  | 214.1157
213.7071 |  | 214.1157
212.526 |  | 214.1158
214.105 |  | 214.1167
211.412 |  | 214.1169
214.9349 |  | 214.1194
212.9999 |  | 214.1201
214.6303 |  | 214.1202
213.2132 |  | 214.121
211.5675 |  | 214.1215
213.5754 |  | 214.1217
213.6707 |  | 214.1223
212.6748 |  | 214.1227
212.8623 |  | 214.1245
212.4224 |  | 214.1254
212.2527 |  | 214.126
213.7094 |  | 214.1265
212.4838 |  | 214.1271
214.8235 |  | 214.1271
214.0038 |  | 214.1271
214.7366 |  | 214.1274
211.9503 |  | 214.1277
213.8018 |  | 214.1278
211.2019 |  | 214.1279
214.6397 |  | 214.1295
212.2767 |  | 214.1308
213.9477 |  | 214.1308
214.8816 |  | 214.1313
214.1539 |  | 214.1316
213.9275 |  | 214.1335
214.4329 |  | 214.1338
211.4034 |  | 214.1349
213.6239 |  | 214.135
212.5272 |  | 214.1362
213.7092 |  | 214.1382
214.7452 |  | 214.1383
212.0761 |  | 214.1388
211.7525 |  | 214.1397
213.7324 |  | 214.1407
211.923 |  | 214.1415
212.0399 |  | 214.1438
212.7612 |  | 214.1439
214.7194 |  | 214.145
212.6369 |  | 214.1459
214.1508 |  | 214.1469
211.9261 |  | 214.147
212.494 |  | 214.1484
211.8651 |  | 214.1488
214.4113 |  | 214.1508
212.9968 |  | 214.1518
214.4749 |  | 214.1524
212.5327 |  | 214.1533
213.8853 |  | 214.1537
213.8786 |  | 214.1539
213.7252 |  | 214.154
214.4833 |  | 214.155
212.1273 |  | 214.1555
212.6826 |  | 214.1558
212.3857 |  | 214.156
213.2373 |  | 214.1562
213.559 |  | 214.1563
213.4453 |  | 214.1574
213.1155 |  | 214.158
212.805 |  | 214.1583
211.9625 |  | 214.1588
210.6363 |  | 214.16
212.9517 |  | 214.1614
213.1905 |  | 214.1615
214.6653 |  | 214.1619
216.2787 |  | 214.162
212.4217 |  | 214.1622
214.2405 |  | 214.1632
213.1641 |  | 214.1653
211.9915 |  | 214.168
214.0095 |  | 214.1682
213.7133 |  | 214.1686
213.4906 |  | 214.1694
214.9204 |  | 214.1704
212.0895 |  | 214.1719
211.9784 |  | 214.1728
212.6335 |  | 214.1777
211.3898 |  | 214.178
213.7765 |  | 214.1782
214.8612 |  | 214.1786
214.8551 |  | 214.1789
213.4728 |  | 214.1792
214.4859 |  | 214.1794
213.661 |  | 214.1796
215.1176 |  | 214.18
215.8313 |  | 214.1801
213.3197 |  | 214.1821
212.7794 |  | 214.1825
213.1998 |  | 214.1825
213.6639 |  | 214.1829
212.5252 |  | 214.1836
213.5677 |  | 214.1836
212.3618 |  | 214.1874
212.2597 |  | 214.1874
213.8283 |  | 214.1876
213.5221 |  | 214.1885
216.7536 |  | 214.1886
213.3586 |  | 214.1892
214.1125 |  | 214.1895
214.0544 |  | 214.1896
213.9958 |  | 214.1908
214.9308 |  | 214.1914
214.0109 |  | 214.1914
213.7763 |  | 214.192
212.078 |  | 214.193
214.3905 |  | 214.1947
213.5329 |  | 214.1949
215.1157 |  | 214.1954
212.854 |  | 214.196
212.0441 |  | 214.1964
211.8975 |  | 214.1968
213.2626 |  | 214.197
214.233 |  | 214.1979
213.8982 |  | 214.1988
213.2447 |  | 214.1988
211.9872 |  | 214.1998
213.8463 |  | 214.2003
212.5219 |  | 214.2007
210.5244 |  | 214.2018
210.2999 |  | 214.2044
214.3205 |  | 214.2059
214.4333 |  | 214.2059
212.8906 |  | 214.2061
214.7469 |  | 214.2063
214.1825 |  | 214.2066
212.9268 |  | 214.2072
214.1157 |  | 214.2076
213.9843 |  | 214.2086
213.2498 |  | 214.2088
212.9397 |  | 214.2093
212.6238 |  | 214.2122
212.7763 |  | 214.2124
214.6091 |  | 214.2135
212.797 |  | 214.214
211.789 |  | 214.2173
214.0166 |  | 214.2178
212.682 |  | 214.219
212.9411 |  | 214.2192
215.1817 |  | 214.2193
215.8607 |  | 214.2206
213.1719 |  | 214.2212
212.2119 |  | 214.2217
212.6743 |  | 214.2233
213.3668 |  | 214.2254
213.1228 |  | 214.2256
214.0497 |  | 214.2265
214.0867 |  | 214.2269
211.9951 |  | 214.2273
215.0737 |  | 214.2277
213.4709 |  | 214.2288
212.392 |  | 214.2292
215.369 |  | 214.2307
215.8821 |  | 214.2311
214.7724 |  | 214.2328
213.4721 |  | 214.233
212.71 |  | 214.2336
213.6533 |  | 214.2342
213.1475 |  | 214.2346
212.8527 |  | 214.2354
214.3367 |  | 214.2355
213.0146 |  | 214.2356
211.9332 |  | 214.2365
213.7417 |  | 214.237
214.4602 |  | 214.2371
212.8466 |  | 214.2384
211.2545 |  | 214.2401
213.0755 |  | 214.2405
211.7356 |  | 214.2411
213.2438 |  | 214.2414
213.4022 |  | 214.2421
215.3318 |  | 214.2427
213.0808 |  | 214.2431
215.432 |  | 214.2433
212.8143 |  | 214.2437
211.9937 |  | 214.2444
214.107 |  | 214.2469
213.8749 |  | 214.2471
213.4328 |  | 214.2475
212.923 |  | 214.2476
213.8375 |  | 214.248
213.0396 |  | 214.2486
215.6067 |  | 214.2511
213.2396 |  | 214.2515
213.9001 |  | 214.2518
213.5381 |  | 214.2519
211.5476 |  | 214.2526
214.2962 |  | 214.2526
212.989 |  | 214.2549
212.4269 |  | 214.255
213.0794 |  | 214.2552
213.7482 |  | 214.2571
212.6867 |  | 214.2594
214.214 |  | 214.2596
213.3888 |  | 214.2598
212.2567 |  | 214.2605
216.0138 |  | 214.2636
214.1227 |  | 214.2643
213.2807 |  | 214.2651
212.8979 |  | 214.2652
216.0188 |  | 214.2657
212.6431 |  | 214.2669
213.0868 |  | 214.2671
212.9285 |  | 214.2684
212.6531 |  | 214.269
211.353 |  | 214.2698
212.6527 |  | 214.27
214.0614 |  | 214.2705
215.0098 |  | 214.2705
211.866 |  | 214.2706
212.175 |  | 214.271
213.5574 |  | 214.2718
212.5336 |  | 214.2748
213.5716 |  | 214.2752
212.1847 |  | 214.2755
213.9358 |  | 214.2762
215.1392 |  | 214.2787
212.7335 |  | 214.2812
213.3256 |  | 214.2813
214.4789 |  | 214.2827
214.5021 |  | 214.2852
214.2973 |  | 214.2863
213.8967 |  | 214.2874
213.5665 |  | 214.2875
211.5752 |  | 214.2878
211.6616 |  | 214.2881
213.9688 |  | 214.2889
211.6463 |  | 214.2902
212.127 |  | 214.2906
213.152 |  | 214.2926
213.7197 |  | 214.2926
214.7603 |  | 214.2932
213.5193 |  | 214.2939
213.1578 |  | 214.2962
213.8269 |  | 214.2973
216.6057 |  | 214.2973
213.0604 |  | 214.2976
212.2436 |  | 214.3012
213.5132 |  | 214.3031
212.4216 |  | 214.3035
213.2425 |  | 214.3037
214.1223 |  | 214.3042
211.5499 |  | 214.3049
212.0097 |  | 214.3051
212.9729 |  | 214.3055
213.6835 |  | 214.3056
212.5207 |  | 214.3069
213.1822 |  | 214.3069
213.8567 |  | 214.3076
213.4499 |  | 214.3085
214.6104 |  | 214.3096
212.8701 |  | 214.3103
214.4213 |  | 214.3106
213.2274 |  | 214.3134
213.3714 |  | 214.3143
213.4279 |  | 214.3149
212.7435 |  | 214.317
214.6625 |  | 214.3177
213.25 |  | 214.3178
214.4531 |  | 214.3178
212.74 |  | 214.3178
213.279 |  | 214.3187
212.6804 |  | 214.3192
215.0168 |  | 214.3204
210.4993 |  | 214.3205
211.7432 |  | 214.3208
212.1602 |  | 214.3216
212.6244 |  | 214.3223
214.9215 |  | 214.3282
213.2635 |  | 214.3286
214.4177 |  | 214.3288
214.0546 |  | 214.3293
212.8564 |  | 214.331
214.7403 |  | 214.3337
213.6741 |  | 214.334
212.6187 |  | 214.3358
214.9972 |  | 214.3366
212.8185 |  | 214.3367
215.0665 |  | 214.3367
214.9216 |  | 214.3379
213.3546 |  | 214.3381
213.0149 |  | 214.3389
213.5616 |  | 214.34
211.4596 |  | 214.3406
213.7066 |  | 214.3407
212.3237 |  | 214.3425
212.9257 |  | 214.3437
213.4781 |  | 214.3454
214.1524 |  | 214.3457
214.0549 |  | 214.3465
213.9127 |  | 214.3479
213.2807 |  | 214.3493
211.5345 |  | 214.3494
214.8132 |  | 214.3496
211.5112 |  | 214.3498
214.6366 |  | 214.35
212.8941 |  | 214.3501
212.1554 |  | 214.3507
214.0296 |  | 214.3516
214.0522 |  | 214.3525
213.5856 |  | 214.3534
212.5233 |  | 214.3569
214.4896 |  | 214.3571
212.9931 |  | 214.3593
214.8547 |  | 214.3604
211.1622 |  | 214.3614
211.8794 |  | 214.3624
214.6005 |  | 214.3625
214.1169 |  | 214.3631
212.1766 |  | 214.3636
211.6791 |  | 214.3666
211.9706 |  | 214.3667
212.4826 |  | 214.3668
213.6739 |  | 214.3672
213.0519 |  | 214.3687
214.2307 |  | 214.369
214.6381 |  | 214.3699
214.1836 |  | 214.3703
214.6194 |  | 214.3703
213.8847 |  | 214.3713
213.7203 |  | 214.3714
211.3405 |  | 214.3728
213.5025 |  | 214.3736
212.5958 |  | 214.3737
212.3847 |  | 214.3749
211.5313 |  | 214.3754
215.401 |  | 214.3757
212.39 |  | 214.3782
213.4467 |  | 214.3801
212.6357 |  | 214.3807
211.6489 |  | 214.3821
212.7686 |  | 214.3828
213.1445 |  | 214.3831
213.1945 |  | 214.3836
211.5145 |  | 214.3877
211.9919 |  | 214.3904
212.3252 |  | 214.3905
213.2709 |  | 214.3944
212.4074 |  | 214.395
213.5966 |  | 214.3956
213.2578 |  | 214.3958
215.655 |  | 214.3978
213.3777 |  | 214.3987
213.6068 |  | 214.3989
213.4307 |  | 214.3991
213.4249 |  | 214.3997
212.0793 |  | 214.3999
212.6684 |  | 214.4009
213.0483 |  | 214.4011
213.2068 |  | 214.4012
213.9294 |  | 214.4039
215.0011 |  | 214.4041
212.6753 |  | 214.4046
216.0382 |  | 214.4058
213.5545 |  | 214.4065
213.5625 |  | 214.407
213.4628 |  | 214.4073
211.7161 |  | 214.4077
212.9254 |  | 214.4079
211.9335 |  | 214.4092
212.7212 |  | 214.4093
214.7181 |  | 214.4104
213.1821 |  | 214.4104
212.691 |  | 214.4113
213.6792 |  | 214.4117
213.8742 |  | 214.413
211.5501 |  | 214.4132
215.0268 |  | 214.4132
213.9361 |  | 214.4177
213.8681 |  | 214.4188
214.0925 |  | 214.4195
213.6149 |  | 214.4213
214.6513 |  | 214.4243
212.0888 |  | 214.4264
212.7501 |  | 214.4275
213.4487 |  | 214.428
212.9682 |  | 214.4307
213.7521 |  | 214.4317
214.1027 |  | 214.4321
214.442 |  | 214.4329
213.8798 |  | 214.4333
214.2526 |  | 214.4335
212.5262 |  | 214.4344
213.8393 |  | 214.4368
212.1487 |  | 214.4383
214.3293 |  | 214.4401
214.1316 |  | 214.4409
212.7578 |  | 214.441
212.7251 |  | 214.441
213.6434 |  | 214.442
212.3998 |  | 214.4438
213.5783 |  | 214.4444
214.1588 |  | 214.4471
213.3367 |  | 214.448
213.9374 |  | 214.4482
214.986 |  | 214.4497
212.4393 |  | 214.4505
214.3728 |  | 214.4514
213.3783 |  | 214.4522
214.1518 |  | 214.4531
212.4615 |  | 214.4539
213.044 |  | 214.4544
213.7205 |  | 214.4549
213.4489 |  | 214.4549
212.724 |  | 214.4553
213.3475 |  | 214.4569
213.6622 |  | 214.4571
212.6459 |  | 214.4576
213.5569 |  | 214.459
213.9908 |  | 214.4598
213.1697 |  | 214.4599
214.3178 |  | 214.4602
212.0399 |  | 214.4602
213.3059 |  | 214.4612
214.3049 |  | 214.4621
213.8296 |  | 214.4623
213.7444 |  | 214.4627
213.9788 |  | 214.463
213.167 |  | 214.4643
213.2907 |  | 214.4644
213.719 |  | 214.4649
213.3298 |  | 214.4651
213.7602 |  | 214.4655
210.5734 |  | 214.466
213.5357 |  | 214.4662
213.9377 |  | 214.4676
211.3084 |  | 214.4693
213.1985 |  | 214.4707
215.5492 |  | 214.4715
211.6737 |  | 214.4724
213.138 |  | 214.4749
211.6647 |  | 214.4752
213.9588 |  | 214.4771
212.9866 |  | 214.4789
213.4059 |  | 214.483
213.829 |  | 214.4831
213.1887 |  | 214.4833
213.7518 |  | 214.4847
214.6049 |  | 214.4848
212.443 |  | 214.4855
213.1474 |  | 214.4859
213.2422 |  | 214.4866
213.6093 |  | 214.4872
213.7729 |  | 214.4879
212.7291 |  | 214.4886
213.0375 |  | 214.4896
214.0256 |  | 214.4898
211.9539 |  | 214.49
215.1325 |  | 214.4903
213.8525 |  | 214.4912
212.5244 |  | 214.4921
214.3958 |  | 214.4932
213.1365 |  | 214.4946
212.3892 |  | 214.4972
214.4317 |  | 214.4972
213.5891 |  | 214.4984
213.7879 |  | 214.4987
213.6208 |  | 214.5002
212.0834 |  | 214.5019
211.3516 |  | 214.5021
212.0141 |  | 214.5046
214.8714 |  | 214.505
212.8812 |  | 214.5079
212.8586 |  | 214.5083
213.2054 |  | 214.5084
213.7628 |  | 214.5088
214.5546 |  | 214.5092
215.0674 |  | 214.5119
212.7386 |  | 214.5122
214.248 |  | 214.5125
213.4458 |  | 214.5167
213.5378 |  | 214.5185
211.8117 |  | 214.5208
211.8471 |  | 214.5211
211.3976 |  | 214.5215
215.863 |  | 214.522
212.4311 |  | 214.522
213.5428 |  | 214.5222
213.6318 |  | 214.5234
215.377 |  | 214.5238
212.2582 |  | 214.5242
214.5208 |  | 214.5291
215.149 |  | 214.5313
212.4284 |  | 214.5339
212.7981 |  | 214.5344
213.2974 |  | 214.5346
215.0688 |  | 214.535
212.577 |  | 214.5373
214.8312 |  | 214.5403
211.5047 |  | 214.5409
213.301 |  | 214.5424
214.5802 |  | 214.5425
212.205 |  | 214.5426
214.0304 |  | 214.5429
213.8788 |  | 214.5429
211.8842 |  | 214.5437
213.7084 |  | 214.5468
212.6303 |  | 214.5478
212.6743 |  | 214.5485
213.547 |  | 214.5546
213.1244 |  | 214.5553
212.682 |  | 214.5556
214.684 |  | 214.5567
214.8031 |  | 214.5572
212.4236 |  | 214.5573
213.49 |  | 214.5576
211.4739 |  | 214.5577
212.9639 |  | 214.559
213.9798 |  | 214.5597
214.317 |  | 214.5598
213.0786 |  | 214.5616
213.2527 |  | 214.5621
213.958 |  | 214.5627
213.5045 |  | 214.564
214.1097 |  | 214.5654
212.2323 |  | 214.5657
212.0721 |  | 214.5662
213.6897 |  | 214.5663
214.4132 |  | 214.5673
213.1394 |  | 214.5674
212.5681 |  | 214.5698
213.3157 |  | 214.571
211.2925 |  | 214.5738
213.6073 |  | 214.5751
212.099 |  | 214.5759
211.1715 |  | 214.5763
213.5924 |  | 214.5764
212.4924 |  | 214.5781
213.081 |  | 214.5785
212.6904 |  | 214.5802
212.4768 |  | 214.5821
212.082 |  | 214.5827
212.8929 |  | 214.583
212.1219 |  | 214.584
214.4009 |  | 214.5853
215.2336 |  | 214.5871
212.7246 |  | 214.5874
211.9087 |  | 214.5882
212.7834 |  | 214.5892
212.3625 |  | 214.5893
214.9717 |  | 214.5925
214.1407 |  | 214.5932
213.4064 |  | 214.5943
215.9795 |  | 214.5946
213.7064 |  | 214.5958
213.3989 |  | 214.5962
213.8563 |  | 214.5968
212.0456 |  | 214.5969
213.4425 |  | 214.5972
212.9447 |  | 214.5972
212.2584 |  | 214.5981
213.5598 |  | 214.5982
214.6496 |  | 214.6005
213.8805 |  | 214.6009
212.8587 |  | 214.6014
216.0297 |  | 214.6024
212.2443 |  | 214.6025
213.5955 |  | 214.6033
214.6615 |  | 214.6037
213.6881 |  | 214.6043
211.6367 |  | 214.6049
213.9202 |  | 214.609
214.5215 |  | 214.6091
214.2371 |  | 214.6104
213.188 |  | 214.6117
213.1266 |  | 214.6127
212.9157 |  | 214.6133
212.0665 |  | 214.6142
212.954 |  | 214.6148
214.7289 |  | 214.6183
214.2748 |  | 214.6194
213.9154 |  | 214.6194
215.1607 |  | 214.6196
213.3879 |  | 214.6207
214.1968 |  | 214.622
212.4897 |  | 214.6224
215.0391 |  | 214.6228
212.774 |  | 214.6233
213.1287 |  | 214.6247
214.3807 |  | 214.6249
212.712 |  | 214.6267
212.7198 |  | 214.6276
212.1584 |  | 214.6282
211.3472 |  | 214.6303
213.5777 |  | 214.6304
214.2651 |  | 214.6328
211.8339 |  | 214.6333
214.7367 |  | 214.6349
213.9734 |  | 214.636
212.9356 |  | 214.6366
215.6434 |  | 214.637
212.3135 |  | 214.6374
212.6501 |  | 214.6381
213.8943 |  | 214.6397
213.6712 |  | 214.6398
212.6289 |  | 214.6398
213.8784 |  | 214.641
214.4012 |  | 214.6447
213.5906 |  | 214.646
213.8433 |  | 214.6464
215.4387 |  | 214.6471
210.8992 |  | 214.6474
213.9968 |  | 214.6492
212.5886 |  | 214.6496
212.6658 |  | 214.6501
213.7407 |  | 214.6501
212.6004 |  | 214.6512
213.5692 |  | 214.6513
215.1852 |  | 214.6524
214.9148 |  | 214.6541
214.6812 |  | 214.6606
212.7143 |  | 214.6615
213.8318 |  | 214.6625
213.8437 |  | 214.6629
211.7956 |  | 214.6653
214.7883 |  | 214.6664
214.0539 |  | 214.6685
213.8739 |  | 214.6687
212.3717 |  | 214.6716
213.229 |  | 214.6729
213.6834 |  | 214.6734
211.8717 |  | 214.6765
214.6792 |  | 214.6766
213.0322 |  | 214.6779
211.8027 |  | 214.6785
214.2718 |  | 214.6792
211.8747 |  | 214.681
209.4289 |  | 214.6812
212.0533 |  | 214.6816
211.4676 |  | 214.6819
211.2532 |  | 214.6823
213.3341 |  | 214.6823
213.4307 |  | 214.6832
212.4482 |  | 214.6833
210.3348 |  | 214.6839
213.3723 |  | 214.684
212.4178 |  | 214.6843
212.6487 |  | 214.6851
212.4294 |  | 214.6862
212.5565 |  | 214.689
212.5177 |  | 214.6891
214.8197 |  | 214.6899
213.659 |  | 214.6907
212.6736 |  | 214.6908
214.0091 |  | 214.6911
213.4234 |  | 214.6929
213.399 |  | 214.693
213.7869 |  | 214.6941
213.5987 |  | 214.6944
211.957 |  | 214.6951
212.3291 |  | 214.6977
213.3067 |  | 214.7001
212.8947 |  | 214.7005
212.7746 |  | 214.7014
212.6998 |  | 214.703
212.6354 |  | 214.7045
214.2926 |  | 214.7059
213.6727 |  | 214.7064
214.1892 |  | 214.7067
212.2915 |  | 214.7079
212.9764 |  | 214.71
213.6522 |  | 214.712
213.0997 |  | 214.716
211.407 |  | 214.7162
211.7455 |  | 214.7173
213.153 |  | 214.7181
213.17 |  | 214.7183
212.0531 |  | 214.7194
211.2206 |  | 214.7222
216.1058 |  | 214.7224
212.5305 |  | 214.7224
211.9755 |  | 214.7227
215.0841 |  | 214.7244
214.015 |  | 214.7289
213.1952 |  | 214.7304
214.7854 |  | 214.731
212.0718 |  | 214.7324
211.7814 |  | 214.7364
215.7602 |  | 214.7366
212.9487 |  | 214.7367
212.4346 |  | 214.7374
213.115 |  | 214.7377
211.7816 |  | 214.7379
212.7938 |  | 214.7388
213.9686 |  | 214.7389
212.2424 |  | 214.7398
209.995 |  | 214.7399
214.0316 |  | 214.7403
213.2411 |  | 214.7411
213.9508 |  | 214.7418
212.5107 |  | 214.7429
213.9692 |  | 214.744
213.6487 |  | 214.7444
213.2399 |  | 214.7452
214.0721 |  | 214.7452
212.9095 |  | 214.7455
212.523 |  | 214.7457
213.1712 |  | 214.7469
213.72 |  | 214.75
211.4804 |  | 214.7503
211.7904 |  | 214.7504
211.5437 |  | 214.7507
214.5759 |  | 214.7511
213.7878 |  | 214.7513
211.3179 |  | 214.7516
212.8316 |  | 214.7517
211.6572 |  | 214.752
213.2572 |  | 214.7521
212.2635 |  | 214.7553
212.7573 |  | 214.7555
212.7299 |  | 214.7573
211.856 |  | 214.7587
213.123 |  | 214.7589
212.0882 |  | 214.7603
214.3736 |  | 214.7631
213.2665 |  | 214.7657
211.6637 |  | 214.7669
213.39 |  | 214.7671
215.0113 |  | 214.768
211.7214 |  | 214.7686
213.0191 |  | 214.7701
213.7011 |  | 214.7701
213.0001 |  | 214.7724
213.375 |  | 214.7725
213.887 |  | 214.7736
213.3981 |  | 214.774
213.1825 |  | 214.775
213.0919 |  | 214.7755
215.1697 |  | 214.7763
213.6733 |  | 214.7778
213.4582 |  | 214.7784
213.403 |  | 214.7789
214.1555 |  | 214.7801
214.2827 |  | 214.7809
213.1345 |  | 214.7835
212.1471 |  | 214.7838
214.0452 |  | 214.7847
213.3753 |  | 214.7854
212.7383 |  | 214.7871
214.3904 |  | 214.7883
213.2012 |  | 214.7892
213.1527 |  | 214.7892
212.9496 |  | 214.7896
212.8679 |  | 214.791
213.6329 |  | 214.7946
214.1362 |  | 214.7979
212.9433 |  | 214.7998
215.6759 |  | 214.8031
213.9693 |  | 214.8033
213.1335 |  | 214.8043
213.8478 |  | 214.8077
212.2768 |  | 214.8087
213.5577 |  | 214.8093
213.3351 |  | 214.8122
213.42 |  | 214.8122
214.8478 |  | 214.8125
211.8639 |  | 214.8132
212.8669 |  | 214.8132
213.7594 |  | 214.8135
214.8803 |  | 214.8143
214.5764 |  | 214.8147
214.6333 |  | 214.8149
214.609 |  | 214.8155
212.9015 |  | 214.8155
215.4928 |  | 214.8197
212.3033 |  | 214.8199
214.0478 |  | 214.8219
213.7418 |  | 214.822
211.9831 |  | 214.823
215.1591 |  | 214.8234
214.5674 |  | 214.8235
213.3885 |  | 214.8236
212.0635 |  | 214.824
213.2831 |  | 214.8243
212.8242 |  | 214.825
214.0047 |  | 214.8253
213.3777 |  | 214.8253
211.7698 |  | 214.828
214.8599 |  | 214.8312
214.7892 |  | 214.834
212.1119 |  | 214.8346
213.4623 |  | 214.8358
213.3179 |  | 214.8407
211.5151 |  | 214.8408
210.6229 |  | 214.8411
214.3493 |  | 214.844
215.4314 |  | 214.8478
212.1691 |  | 214.8484
211.9031 |  | 214.8543
214.8155 |  | 214.8545
213.2122 |  | 214.8547
214.0627 |  | 214.8551
214.156 |  | 214.8553
212.6801 |  | 214.8561
212.647 |  | 214.857
212.0694 |  | 214.8571
213.4749 |  | 214.8596
214.7838 |  | 214.8599
212.3998 |  | 214.8612
213.1447 |  | 214.8617
211.805 |  | 214.8625
214.6224 |  | 214.8631
211.9699 |  | 214.8635
213.2959 |  | 214.8714
212.4699 |  | 214.874
213.7681 |  | 214.8747
213.9789 |  | 214.8782
212.5891 |  | 214.8798
213.8121 |  | 214.8802
214.16 |  | 214.8803
212.5868 |  | 214.8811
213.4573 |  | 214.8816
211.33 |  | 214.8851
214.1004 |  | 214.8879
212.8347 |  | 214.8887
212.4265 |  | 214.8887
213.8762 |  | 214.8888
214.7045 |  | 214.8889
213.503 |  | 214.8954
211.852 |  | 214.8956
213.2276 |  | 214.9023
213.0369 |  | 214.9148
213.0394 |  | 214.9148
214.8143 |  | 214.9149
213.1599 |  | 214.9157
214.8234 |  | 214.9171
211.9756 |  | 214.9182
214.2906 |  | 214.9204
214.0861 |  | 214.921
213.4796 |  | 214.9215
212.3842 |  | 214.9216
213.5318 |  | 214.9228
214.4401 |  | 214.9232
212.9475 |  | 214.9281
212.3444 |  | 214.9288
212.3934 |  | 214.9308
213.3915 |  | 214.9315
212.2279 |  | 214.9321
214.8253 |  | 214.9328
212.8209 |  | 214.934
213.2886 |  | 214.9341
212.178 |  | 214.9349
213.6442 |  | 214.9359
212.8667 |  | 214.9379
213.6956 |  | 214.9386
216.4837 |  | 214.9424
212.9702 |  | 214.9432
216.2387 |  | 214.945
216.1459 |  | 214.9452
215.2292 |  | 214.946
211.1193 |  | 214.9483
211.891 |  | 214.9531
213.6468 |  | 214.9534
213.6228 |  | 214.9539
213.2183 |  | 214.9568
213.5647 |  | 214.9576
213.665 |  | 214.963
210.9027 |  | 214.9642
212.8483 |  | 214.9707
212.4344 |  | 214.9712
212.5206 |  | 214.9717
213.5479 |  | 214.9737
213.3109 |  | 214.9758
214.2549 |  | 214.9763
212.3125 |  | 214.9802
214.3187 |  | 214.9813
211.1124 |  | 214.9835
213.4942 |  | 214.9836
213.5829 |  | 214.9836
213.237 |  | 214.9845
213.2546 |  | 214.986
213.735 |  | 214.9932
213.5659 |  | 214.9934
211.3648 |  | 214.9952
213.0141 |  | 214.9971
212.7065 |  | 214.9972
214.7399 |  | 214.9997
214.3192 |  | 215.0001
212.4503 |  | 215.0011
212.9768 |  | 215.0014
213.597 |  | 215.0017
214.9432 |  | 215.0028
212.4145 |  | 215.0047
213.0425 |  | 215.0057
212.3367 |  | 215.0058
212.6956 |  | 215.0068
213.5029 |  | 215.0081
213.4321 |  | 215.0098
211.8441 |  | 215.0113
213.257 |  | 215.0152
213.5405 |  | 215.0168
214.0665 |  | 215.02
212.2923 |  | 215.0206
211.6848 |  | 215.0225
212.8252 |  | 215.0248
213.6115 |  | 215.0254
213.2934 |  | 215.0268
214.0244 |  | 215.0278
215.7676 |  | 215.029
211.8553 |  | 215.0296
213.675 |  | 215.0324
214.0426 |  | 215.0342
213.557 |  | 215.0349
214.3479 |  | 215.0357
215.3258 |  | 215.0368
212.1731 |  | 215.0391
213.4773 |  | 215.0472
212.224 |  | 215.0502
212.9501 |  | 215.0622
214.5982 |  | 215.063
212.7066 |  | 215.0665
213.2446 |  | 215.0674
212.633 |  | 215.0681
212.166 |  | 215.0685
212.0256 |  | 215.0688
212.8284 |  | 215.0737
212.4472 |  | 215.0766
214.35 |  | 215.079
214.9952 |  | 215.08
212.1848 |  | 215.0819
211.2004 |  | 215.0835
211.6044 |  | 215.0841
211.9722 |  | 215.0845
212.9758 |  | 215.0893
213.3905 |  | 215.0902
212.8729 |  | 215.0903
212.0836 |  | 215.0944
213.3403 |  | 215.0967
213.5271 |  | 215.0976
212.5651 |  | 215.0991
213.8216 |  | 215.0995
214.5079 |  | 215.1085
212.6868 |  | 215.1119
211.452 |  | 215.112
213.4996 |  | 215.1157
213.4716 |  | 215.1176
213.7733 |  | 215.118
214.1908 |  | 215.1181
213.9396 |  | 215.1196
212.7833 |  | 215.123
211.5273 |  | 215.1278
213.3351 |  | 215.1325
213.9332 |  | 215.1392
213.4846 |  | 215.1403
212.4848 |  | 215.1404
212.1479 |  | 215.142
212.6908 |  | 215.1426
211.5087 |  | 215.1428
214.1979 |  | 215.1434
211.3886 |  | 215.148
215.2416 |  | 215.149
213.039 |  | 215.155
211.9501 |  | 215.1555
211.7041 |  | 215.1555
212.8737 |  | 215.1556
213.9041 |  | 215.1584
211.126 |  | 215.1591
214.9379 |  | 215.1601
213.2575 |  | 215.1607
213.8858 |  | 215.161
214.6233 |  | 215.1697
213.4112 |  | 215.1731
214.8087 |  | 215.1813
213.5039 |  | 215.1817
213.9023 |  | 215.1836
212.2482 |  | 215.1852
214.4444 |  | 215.1877
213.0487 |  | 215.1919
212.7381 |  | 215.195
214.3593 |  | 215.2059
213.559 |  | 215.2081
213.0522 |  | 215.2203
212.8618 |  | 215.222
212.2852 |  | 215.2235
214.0563 |  | 215.2261
213.5571 |  | 215.2266
213.0215 |  | 215.2292
212.7134 |  | 215.2329
213.8534 |  | 215.2336
213.5956 |  | 215.2352
212.5024 |  | 215.2369
213.5271 |  | 215.2377
214.4655 |  | 215.2385
211.4467 |  | 215.2416
212.7657 |  | 215.2442
212.8927 |  | 215.2453
212.5708 |  | 215.2459
212.1221 |  | 215.2481
213.7014 |  | 215.2488
213.7437 |  | 215.2546
212.4151 |  | 215.2558
213.7321 |  | 215.2571
214.2273 |  | 215.2614
214.5663 |  | 215.265
213.0611 |  | 215.2706
213.0348 |  | 215.2804
212.4525 |  | 215.2811
213.8665 |  | 215.2829
212.1361 |  | 215.3036
211.3018 |  | 215.309
213.0594 |  | 215.3095
213.7037 |  | 215.3173
213.1162 |  | 215.3221
213.7108 |  | 215.3258
213.353 |  | 215.3271
214.5429 |  | 215.3278
211.0322 |  | 215.3311
213.4971 |  | 215.3318
213.4519 |  | 215.3324
211.9449 |  | 215.3355
214.6194 |  | 215.3359
213.4613 |  | 215.343
213.1069 |  | 215.3453
214.2605 |  | 215.3462
212.4611 |  | 215.3489
213.4025 |  | 215.3508
212.2566 |  | 215.3603
213.6274 |  | 215.3617
213.2359 |  | 215.3645
213.7639 |  | 215.3648
216.8752 |  | 215.369
213.0889 |  | 215.3729
212.777 |  | 215.3734
210.8258 |  | 215.377
213.0315 |  | 215.3774
211.2237 |  | 215.3788
211.721 |  | 215.382
214.8149 |  | 215.3892
214.6304 |  | 215.391
212.4376 |  | 215.3924
215.5702 |  | 215.3939
213.2365 |  | 215.394
212.8991 |  | 215.401
213.829 |  | 215.4049
213.476 |  | 215.4144
210.2318 |  | 215.4163
213.3266 |  | 215.4223
214.5185 |  | 215.427
212.9561 |  | 215.4276
213.169 |  | 215.4295
214.5485 |  | 215.4314
212.4424 |  | 215.432
212.6222 |  | 215.4387
211.5674 |  | 215.4397
212.8088 |  | 215.4477
213.4983 |  | 215.4484
211.4541 |  | 215.4552
211.718 |  | 215.4561
214.681 |  | 215.4585
213.913 |  | 215.4595
210.7163 |  | 215.4602
215.7935 |  | 215.476
212.8666 |  | 215.477
212.3399 |  | 215.479
213.9238 |  | 215.4888
211.6063 |  | 215.4928
213.2949 |  | 215.4977
213.9534 |  | 215.498
212.1133 |  | 215.5009
214.395 |  | 215.5114
212.1806 |  | 215.518
212.6082 |  | 215.5221
212.9302 |  | 215.5302
213.2066 |  | 215.5381
212.764 |  | 215.5492
212.8031 |  | 215.5494
212.5612 |  | 215.5504
212.9639 |  | 215.5608
214.1201 |  | 215.563
212.0462 |  | 215.5674
213.0817 |  | 215.5675
213.2528 |  | 215.5702
213.8824 |  | 215.5723
215.7908 |  | 215.5912
212.9638 |  | 215.5942
213.9031 |  | 215.6064
214.9813 |  | 215.6067
212.9948 |  | 215.6079
212.7221 |  | 215.6409
211.4185 |  | 215.6434
214.0947 |  | 215.644
213.2924 |  | 215.6443
214.0553 |  | 215.6448
213.0108 |  | 215.655
213.9244 |  | 215.656
212.2351 |  | 215.6575
212.1333 |  | 215.6617
214.5853 |  | 215.6667
214.4505 |  | 215.6686
214.703 |  | 215.6717
214.7001 |  | 215.6741
213.8508 |  | 215.6759
215.5674 |  | 215.6792
213.7489 |  | 215.6802
215.825 |  | 215.6804
212.5483 |  | 215.6845
214.2932 |  | 215.6976
212.8511 |  | 215.6993
212.3312 |  | 215.7002
211.6868 |  | 215.7025
212.8963 |  | 215.7154
214.6148 |  | 215.7193
213.5266 |  | 215.7204
213.56 |  | 215.7216
213.4484 |  | 215.7288
213.0284 |  | 215.7371
213.3369 |  | 215.7457
214.1122 |  | 215.7529
213.0188 |  | 215.7602
214.1836 |  | 215.7676
212.9804 |  | 215.7818
211.6297 |  | 215.7908
213.6226 |  | 215.7935
214.154 |  | 215.7949
212.3078 |  | 215.8012
214.4104 |  | 215.8044
212.9278 |  | 215.8174
214.0103 |  | 215.8188
213.1008 |  | 215.825
213.0805 |  | 215.8313
213.4638 |  | 215.8389
214.2288 |  | 215.8607
212.4843 |  | 215.863
214.3494 |  | 215.8731
212.0883 |  | 215.8778
211.0471 |  | 215.8787
212.5871 |  | 215.8821
212.7912 |  | 215.884
213.9107 |  | 215.8912
211.1092 |  | 215.8914
212.9811 |  | 215.9002
212.7983 |  | 215.9218
214.334 |  | 215.9239
213.7042 |  | 215.9572
213.3476 |  | 215.9646
213.5011 |  | 215.9795
212.5202 |  | 216.0092
211.961 |  | 216.0138
213.5349 |  | 216.0188
212.0933 |  | 216.0209
211.7073 |  | 216.0297
212.897 |  | 216.0382
214.0907 |  | 216.0576
213.0287 |  | 216.0657
216.7143 |  | 216.1058
213.0748 |  | 216.1345
212.7054 |  | 216.1368
213.8854 |  | 216.141
212.6874 |  | 216.1459
213.1869 |  | 216.1854
213.3358 |  | 216.208
215.394 |  | 216.2387
214.6929 |  | 216.2553
212.1226 |  | 216.2787
212.8976 |  | 216.3374
213.3486 |  | 216.342
213.1997 |  | 216.3815
213.7896 |  | 216.3867
213.746 |  | 216.4233
213.8346 |  | 216.4317
213.3099 |  | 216.4837
212.9154 |  | 216.5399
213.8017 |  | 216.5549
213.3592 |  | 216.6057
213.8541 |  | 216.6535
215.0057 |  | 216.7143
213.6383 |  | 216.7536
213.6349 |  | 216.7765
215.1404 |  | 216.778
213.3676 |  | 216.837
212.4715 |  | 216.8752
213.2005 |  | 216.9912
214.1949 |  | 217.1